MediPi - NHS Digital Functional Test Schedule and Evidence

[1 Homepage - General Screen 11](#_Toc256000000)

[2 Tile Configuration 11](#_Toc256000001)

[3 Logon - Admin Logon (Called Test Only) 16](#_Toc256000002)

[4 Logon - General and Patient Logon 21](#_Toc256000003)

[5 Logon - MediPi Session Lock 23](#_Toc256000004)

[6 Admin Configuration Screen 27](#_Toc256000005)

[7 Finger Oximeter - General Screen 33](#_Toc256000006)

[8 Finger Oximeter - Record 35](#_Toc256000007)

[9 Finger Oximeter - Transmit 37](#_Toc256000008)

[10 Thermometer - General Screen 39](#_Toc256000009)

[11 Thermometer - Record 42](#_Toc256000010)

[12 Thermometer - Transmit 44](#_Toc256000011)

[13 Diagnostic Scale - General Screen 46](#_Toc256000012)

[14 Diagnostic Scale - Record 48](#_Toc256000013)

[15 Diagnostic Scale - Transmit 50](#_Toc256000014)

[16 Blood Pressure - General Screen 52](#_Toc256000015)

[17 Blood Pressure - Record 54](#_Toc256000016)

[18 Blood Pressure - Transmit 56](#_Toc256000017)

[19 COPD Questionnaire 58](#_Toc256000018)

[20 COPD Questionnaire - Transmit 59](#_Toc256000019)

[21 COPD Questionnaire Alerts 60](#_Toc256000020)

[22 Diabetes Questionnaire 61](#_Toc256000021)

[23 Diabetes Questionnaire - Transmit 63](#_Toc256000022)

[24 Diabetes Questionnaire Alerts 64](#_Toc256000023)

[25 Heart Failure Questionnaire 65](#_Toc256000024)

[26 Heart Questionnaire - Alerts 67](#_Toc256000025)

[27 Heart Questionnaire - Transmit 68](#_Toc256000026)

[28 Scheduler - Schedule Screens 69](#_Toc256000027)

[29 Scheduler - Transmit (All Data) 76](#_Toc256000028)

[30 Scheduler - Transmit (Individual) 80](#_Toc256000029)

[31 Scheduler - Transmit Summary Screen 87](#_Toc256000030)

[32 Scheduler Configuration 90](#_Toc256000031)

[33 Clinician Messages - General Screen 93](#_Toc256000032)

[34 Clinician Messages - Polling Warning 96](#_Toc256000033)

[35 Transmitter - General Screen 97](#_Toc256000034)

[36 Transmitter - Transmit (All Data) 100](#_Toc256000035)

[37 Logon & Logoff 105](#_Toc256000036)

[38 Session Timeout 107](#_Toc256000037)

[39 Caseload View - Exclamation Mark 109](#_Toc256000038)

[40 Caseload View - Happy Face 111](#_Toc256000039)

[41 Caseload View - Question Mark 114](#_Toc256000040)

[42 Caseload View - Question Mark Reset 115](#_Toc256000041)

[43 Caseload View - Sad Face 118](#_Toc256000042)

[44 Browser Testing - Internet Explorer 120](#_Toc256000043)

[45 Homepage - General Screen 121](#_Toc256000044)

[46 Patient Detail View 122](#_Toc256000045)

[47 Patients Caseload View 126](#_Toc256000046)

[48 Data transmitted from multiple MediPi Devices 129](#_Toc256000047)

[49 Fatal Error Screen 132](#_Toc256000048)

[50 Loss of WIFI Connection 132](#_Toc256000049)

[51 MediPi Patient Banner - Patient Name 135](#_Toc256000050)

[52 MediPi Patient Banner (Called Test Only) 136](#_Toc256000051)

[53 Multiple Connected Devices 138](#_Toc256000052)

[54 MediPi Application - Order of readings changed via schedule 139](#_Toc256000053)

[55 MediPi Application - Removal of all tiles on the home screen except for the readings tile 140](#_Toc256000054)

[56 MediPi Application - Removal of PID on main MediPi Screen before authentication 142](#_Toc256000055)

[57 MediPi Application - Timeout increased to 15 minutes 143](#_Toc256000056)

[58 MediPi Clinical Application - Weight thresholds changed to be simple high & low 145](#_Toc256000057)

[59 Blood Pressure Threshold 148](#_Toc256000058)

[60 COPD Questionnaire Results 154](#_Toc256000059)

[61 Diabetes Questionnaire Results 156](#_Toc256000060)

[62 Heart Questionnaire Results 157](#_Toc256000061)

[63 Oxygen Saturation Threshold 159](#_Toc256000062)

[64 Pulse Rate Threshold 165](#_Toc256000063)

[65 Temperature Threshold 171](#_Toc256000064)

[66 Threshold Input Value Validation 176](#_Toc256000065)

[67 Weight Threshold 185](#_Toc256000066)

[68 Blood Pressure Monitor - Time Synchronisation 191](#_Toc256000067)

[69 Finger Oximeter - Time Synchronisation 192](#_Toc256000068)

[70 MediPi Time Synchronisation 194](#_Toc256000069)

[71 Runs 198](#_Toc256000070)

[71.1 Run ID : 28741 - Fast\_Run\_10-19\_10-58-32 198](#_Toc256000071)

[71.2 Run ID : 28740 - Fast\_Run\_10-19\_10-58-28 198](#_Toc256000072)

[71.3 Run ID : 28739 - Fast\_Run\_10-19\_10-58-25 199](#_Toc256000073)

[71.4 Run ID : 28738 - Fast\_Run\_10-19\_10-58-23 199](#_Toc256000074)

[71.5 Run ID : 28737 - Fast\_Run\_10-19\_10-58-20 200](#_Toc256000075)

[71.6 Run ID : 28736 - Fast\_Run\_10-19\_10-58-16 200](#_Toc256000076)

[71.7 Run ID : 28735 - Fast\_Run\_10-19\_10-58-14 201](#_Toc256000077)

[71.8 Run ID : 28734 - Fast\_Run\_10-19\_10-58-4 201](#_Toc256000078)

[71.9 Run ID : 28733 - Fast\_Run\_10-19\_10-58-0 202](#_Toc256000079)

[71.10 Run ID : 28728 - Fast\_Run\_9-21\_9-40-47 202](#_Toc256000080)

[71.11 Run ID : 28727 - Fast\_Run\_9-20\_15-12-26 203](#_Toc256000081)

[71.12 Run ID : 28726 - Fast\_Run\_9-20\_15-12-18 203](#_Toc256000082)

[71.13 Run ID : 28725 - Fast\_Run\_9-20\_15-12-17 204](#_Toc256000083)

[71.14 Run ID : 28724 - Fast\_Run\_9-20\_15-12-15 204](#_Toc256000084)

[71.15 Run ID : 28723 - Fast\_Run\_9-20\_15-12-1 205](#_Toc256000085)

[71.16 Run ID : 28722 - Fast\_Run\_9-20\_15-12-0 205](#_Toc256000086)

[71.17 Run ID : 28721 - Fast\_Run\_9-20\_15-11-59 206](#_Toc256000087)

[71.18 Run ID : 28720 - Fast\_Run\_9-20\_15-11-58 206](#_Toc256000088)

[71.19 Run ID : 28719 - Fast\_Run\_9-20\_15-11-57 207](#_Toc256000089)

[71.20 Run ID : 28718 - Fast\_Run\_9-20\_15-11-56 207](#_Toc256000090)

[71.21 Run ID : 28717 - Fast\_Run\_9-20\_15-11-54 208](#_Toc256000091)

[71.22 Run ID : 28716 - Fast\_Run\_9-20\_15-11-35 208](#_Toc256000092)

[71.23 Run ID : 28715 - Fast\_Run\_9-20\_15-11-21 209](#_Toc256000093)

[71.24 Run ID : 28714 - Fast\_Run\_9-20\_15-11-19 209](#_Toc256000094)

[71.25 Run ID : 28713 - Fast\_Run\_9-20\_15-11-17 210](#_Toc256000095)

[71.26 Run ID : 28712 - Fast\_Run\_9-20\_15-11-16 210](#_Toc256000096)

[71.27 Run ID : 28711 - Fast\_Run\_9-20\_15-11-14 211](#_Toc256000097)

[71.28 Run ID : 28710 - Fast\_Run\_9-20\_15-11-13 211](#_Toc256000098)

[71.29 Run ID : 28709 - Fast\_Run\_9-20\_15-10-52 212](#_Toc256000099)

[71.30 Run ID : 28708 - Fast\_Run\_9-20\_15-10-48 212](#_Toc256000100)

[71.31 Run ID : 28707 - Fast\_Run\_9-20\_15-10-46 213](#_Toc256000101)

[71.32 Run ID : 28706 - Fast\_Run\_9-20\_15-10-43 213](#_Toc256000102)

[71.33 Run ID : 28535 - Run\_7-25\_12-48-27 214](#_Toc256000103)

[71.34 Run ID : 28529 - Run\_7-24\_17-56-45 214](#_Toc256000104)

[71.35 Run ID : 28528 - Run\_7-24\_17-47-7 215](#_Toc256000105)

[71.36 Run ID : 28527 - Run\_7-24\_17-30-33 215](#_Toc256000106)

[71.37 Run ID : 28526 - Run\_7-24\_17-19-12 216](#_Toc256000107)

[71.38 Run ID : 28525 - Run\_7-24\_17-13-54 216](#_Toc256000108)

[71.39 Run ID : 28523 - Run\_7-24\_17-13-38 217](#_Toc256000109)

[71.40 Run ID : 28522 - Run\_7-24\_17-13-35 217](#_Toc256000110)

[71.41 Run ID : 28521 - Run\_7-24\_17-10-22 217](#_Toc256000111)

[71.42 Run ID : 28519 - Run\_7-24\_17-8-25 218](#_Toc256000112)

[71.43 Run ID : 28163 - Run\_6-21\_14-19-35 218](#_Toc256000113)

[71.44 Run ID : 28162 - Run\_6-21\_14-19-24 219](#_Toc256000114)

[71.45 Run ID : 28161 - Run\_6-21\_14-18-6 219](#_Toc256000115)

[71.46 Run ID : 28160 - Run\_6-21\_14-17-54 220](#_Toc256000116)

[71.47 Run ID : 28159 - Run\_6-21\_14-17-27 221](#_Toc256000117)

[71.48 Run ID : 28055 - Fast\_Run\_5-9\_12-55-39 221](#_Toc256000118)

[71.49 Run ID : 28054 - Fast\_Run\_5-9\_12-55-36 222](#_Toc256000119)

[71.50 Run ID : 28053 - Fast\_Run\_5-9\_12-55-36 222](#_Toc256000120)

[71.51 Run ID : 28052 - Fast\_Run\_5-9\_12-55-35 223](#_Toc256000121)

[71.52 Run ID : 28051 - Fast\_Run\_5-9\_12-55-34 223](#_Toc256000122)

[71.53 Run ID : 28050 - Fast\_Run\_5-9\_12-55-33 224](#_Toc256000123)

[71.54 Run ID : 28049 - Fast\_Run\_5-9\_12-55-33 224](#_Toc256000124)

[71.55 Run ID : 28048 - Fast\_Run\_5-9\_12-55-32 225](#_Toc256000125)

[71.56 Run ID : 28047 - Fast\_Run\_5-9\_12-55-31 225](#_Toc256000126)

[71.57 Run ID : 28046 - Fast\_Run\_5-9\_12-55-26 226](#_Toc256000127)

[71.58 Run ID : 28045 - Fast\_Run\_5-9\_12-55-23 226](#_Toc256000128)

[71.59 Run ID : 28044 - Fast\_Run\_5-9\_12-55-21 227](#_Toc256000129)

[71.60 Run ID : 28043 - Fast\_Run\_5-9\_12-55-20 227](#_Toc256000130)

[71.61 Run ID : 28042 - Fast\_Run\_5-9\_12-55-18 227](#_Toc256000131)

[71.62 Run ID : 28041 - Fast\_Run\_5-9\_12-55-17 228](#_Toc256000132)

[71.63 Run ID : 28040 - Fast\_Run\_5-9\_12-55-16 228](#_Toc256000133)

[71.64 Run ID : 28039 - Fast\_Run\_5-9\_12-55-12 229](#_Toc256000134)

[71.65 Run ID : 28038 - Fast\_Run\_5-9\_12-55-8 229](#_Toc256000135)

[71.66 Run ID : 28037 - Fast\_Run\_5-9\_12-55-6 230](#_Toc256000136)

[71.67 Run ID : 28036 - Fast\_Run\_5-9\_12-55-3 230](#_Toc256000137)

[71.68 Run ID : 28035 - Fast\_Run\_5-9\_12-54-59 231](#_Toc256000138)

[71.69 Run ID : 28034 - Fast\_Run\_5-9\_12-54-55 231](#_Toc256000139)

[71.70 Run ID : 28033 - Fast\_Run\_5-9\_12-54-51 232](#_Toc256000140)

[71.71 Run ID : 28032 - Fast\_Run\_5-9\_12-54-50 232](#_Toc256000141)

[71.72 Run ID : 28031 - Fast\_Run\_5-9\_12-54-48 233](#_Toc256000142)

[71.73 Run ID : 28030 - Fast\_Run\_5-9\_12-54-47 233](#_Toc256000143)

[71.74 Run ID : 28029 - Fast\_Run\_5-9\_12-54-46 234](#_Toc256000144)

[71.75 Run ID : 28028 - Fast\_Run\_5-9\_12-54-46 234](#_Toc256000145)

[71.76 Run ID : 28027 - Fast\_Run\_5-9\_12-54-45 235](#_Toc256000146)

[71.77 Run ID : 28026 - Fast\_Run\_5-9\_12-54-43 235](#_Toc256000147)

[71.78 Run ID : 28025 - Fast\_Run\_5-9\_12-54-42 235](#_Toc256000148)

[71.79 Run ID : 28024 - Fast\_Run\_5-9\_12-54-36 236](#_Toc256000149)

[71.80 Run ID : 28023 - Fast\_Run\_5-9\_12-54-33 236](#_Toc256000150)

[71.81 Run ID : 28022 - Fast\_Run\_5-9\_12-54-30 237](#_Toc256000151)

[71.82 Run ID : 28021 - Run\_5-9\_11-13-28 237](#_Toc256000152)

[71.83 Run ID : 28020 - Run\_5-9\_11-12-50 238](#_Toc256000153)

[71.84 Run ID : 27844 - Fast\_Run\_4-3\_15-20-20 238](#_Toc256000154)

[71.85 Run ID : 27843 - Run\_4-3\_14-46-28 239](#_Toc256000155)

[71.86 Run ID : 27842 - Run\_4-3\_14-45-59 239](#_Toc256000156)

[71.87 Run ID : 27841 - Run\_4-3\_14-29-50 240](#_Toc256000157)

[71.88 Run ID : 27840 - Run\_4-3\_14-21-55 240](#_Toc256000158)

[71.89 Run ID : 27838 - Fast\_Run\_4-3\_14-16-50 241](#_Toc256000159)

[71.90 Run ID : 27713 - Run\_3-23\_10-40-17 241](#_Toc256000160)

[71.91 Run ID : 27712 - Fast\_Run\_3-23\_10-36-41 242](#_Toc256000161)

[71.92 Run ID : 27711 - Run\_3-23\_10-36-29 242](#_Toc256000162)

[71.93 Run ID : 27710 - Fast\_Run\_3-23\_10-36-28 243](#_Toc256000163)

[71.94 Run ID : 27709 - Run\_3-23\_10-36-16 243](#_Toc256000164)

[71.95 Run ID : 27708 - Fast\_Run\_3-23\_10-36-15 244](#_Toc256000165)

[71.96 Run ID : 27707 - Run\_3-23\_10-34-23 244](#_Toc256000166)

[71.97 Run ID : 27706 - Run\_3-23\_10-32-23 245](#_Toc256000167)

[71.98 Run ID : 27705 - Run\_3-22\_13-55-54 245](#_Toc256000168)

[71.99 Run ID : 27704 - Run\_3-22\_13-53-18 245](#_Toc256000169)

[71.100 Run ID : 27703 - Run\_3-22\_13-52-41 246](#_Toc256000170)

[71.101 Run ID : 27702 - Run\_3-22\_12-19-56 246](#_Toc256000171)

[71.102 Run ID : 27701 - Run\_3-22\_11-45-52 247](#_Toc256000172)

[71.103 Run ID : 27700 - Run\_3-22\_11-14-14 247](#_Toc256000173)

[71.104 Run ID : 27699 - Run\_3-22\_10-11-39 248](#_Toc256000174)

[71.105 Run ID : 27698 - Run\_3-22\_10-8-54 248](#_Toc256000175)

[71.106 Run ID : 27697 - Run\_3-22\_10-3-12 249](#_Toc256000176)

[71.107 Run ID : 27696 - Run\_3-22\_10-1-20 249](#_Toc256000177)

[71.108 Run ID : 27695 - Run\_3-22\_9-40-49 250](#_Toc256000178)

[71.109 Run ID : 27694 - Run\_3-22\_9-23-52 250](#_Toc256000179)

[71.110 Run ID : 27693 - Run\_3-22\_9-17-3 251](#_Toc256000180)

[71.111 Run ID : 27692 - Run\_3-22\_9-15-3 251](#_Toc256000181)

[71.112 Run ID : 27690 - Run\_3-21\_15-32-46 252](#_Toc256000182)

[71.113 Run ID : 27689 - Run\_3-21\_15-29-19 252](#_Toc256000183)

[71.114 Run ID : 27688 - Run\_3-21\_15-26-14 253](#_Toc256000184)

[71.115 Run ID : 27687 - Run\_3-21\_15-11-23 253](#_Toc256000185)

[71.116 Run ID : 27686 - Fast\_Run\_3-21\_15-11-22 254](#_Toc256000186)

[71.117 Run ID : 27685 - Run\_3-21\_15-10-15 254](#_Toc256000187)

[71.118 Run ID : 27684 - Fast\_Run\_3-21\_15-10-12 255](#_Toc256000188)

[71.119 Run ID : 27683 - Run\_3-21\_15-8-12 255](#_Toc256000189)

[71.120 Run ID : 27682 - Run\_3-21\_15-3-49 255](#_Toc256000190)

[71.121 Run ID : 27681 - Run\_3-21\_15-2-48 256](#_Toc256000191)

[71.122 Run ID : 27680 - Run\_3-21\_15-2-0 256](#_Toc256000192)

[71.123 Run ID : 27679 - Run\_3-21\_15-1-46 257](#_Toc256000193)

[71.124 Run ID : 27678 - Run\_3-21\_15-0-28 257](#_Toc256000194)

[71.125 Run ID : 27677 - Run\_3-21\_15-0-2 258](#_Toc256000195)

[71.126 Run ID : 27676 - Run\_3-21\_14-58-47 258](#_Toc256000196)

[71.127 Run ID : 27675 - Run\_3-21\_14-56-15 259](#_Toc256000197)

[71.128 Run ID : 27674 - Run\_3-21\_14-55-21 259](#_Toc256000198)

[71.129 Run ID : 27673 - Run\_3-21\_14-43-59 260](#_Toc256000199)

[71.130 Run ID : 27672 - Run\_3-21\_14-42-36 260](#_Toc256000200)

[71.131 Run ID : 27671 - Run\_3-21\_14-41-4 261](#_Toc256000201)

[71.132 Run ID : 27670 - Run\_3-21\_14-33-40 261](#_Toc256000202)

[71.133 Run ID : 27669 - Run\_3-21\_14-31-47 262](#_Toc256000203)

[71.134 Run ID : 27668 - Run\_3-21\_14-29-49 262](#_Toc256000204)

[71.135 Run ID : 27667 - Run\_3-21\_14-25-28 263](#_Toc256000205)

[71.136 Run ID : 27666 - Run\_3-21\_14-22-53 263](#_Toc256000206)

[71.137 Run ID : 27665 - Run\_3-21\_13-50-28 264](#_Toc256000207)

[71.138 Run ID : 27664 - Fast\_Run\_3-21\_13-21-19 264](#_Toc256000208)

[71.139 Run ID : 27663 - Fast\_Run\_3-21\_13-21-16 265](#_Toc256000209)

[71.140 Run ID : 27662 - Fast\_Run\_3-21\_13-21-12 265](#_Toc256000210)

[71.141 Run ID : 27661 - Fast\_Run\_3-21\_13-21-9 265](#_Toc256000211)

[71.142 Run ID : 27660 - Fast\_Run\_3-21\_13-21-7 266](#_Toc256000212)

[71.143 Run ID : 27659 - Fast\_Run\_3-21\_13-21-4 266](#_Toc256000213)

[71.144 Run ID : 27656 - Run\_3-21\_11-51-5 267](#_Toc256000214)

[71.145 Run ID : 27655 - Run\_3-21\_11-49-29 267](#_Toc256000215)

[71.146 Run ID : 27653 - Run\_3-21\_11-39-38 268](#_Toc256000216)

[71.147 Run ID : 27652 - Run\_3-21\_11-33-9 268](#_Toc256000217)

[71.148 Run ID : 27651 - Run\_3-21\_11-25-43 269](#_Toc256000218)

[71.149 Run ID : 27650 - Run\_3-21\_11-22-48 269](#_Toc256000219)

[71.150 Run ID : 27649 - Run\_3-21\_11-16-32 270](#_Toc256000220)

[71.151 Run ID : 27648 - Run\_3-21\_11-10-1 270](#_Toc256000221)

[71.152 Run ID : 27647 - Run\_3-21\_11-7-55 271](#_Toc256000222)

[71.153 Run ID : 27646 - Run\_3-21\_11-2-59 271](#_Toc256000223)

[71.154 Run ID : 27644 - Run\_3-21\_10-56-19 272](#_Toc256000224)

[71.155 Run ID : 27105 - Run\_1-27\_14-22-3 272](#_Toc256000225)

[71.156 Run ID : 27104 - Run\_1-27\_11-34-28 273](#_Toc256000226)

[71.157 Run ID : 27103 - Run\_1-27\_11-14-6 273](#_Toc256000227)

[71.158 Run ID : 27102 - Fast\_Run\_1-26\_12-42-49 274](#_Toc256000228)

[71.159 Run ID : 27101 - Run\_1-26\_12-42-12 274](#_Toc256000229)

[71.160 Run ID : 27100 - Run\_1-26\_12-42-1 275](#_Toc256000230)

[71.161 Run ID : 27099 - Run\_1-26\_12-41-47 275](#_Toc256000231)

[71.162 Run ID : 27098 - Run\_1-26\_12-11-47 275](#_Toc256000232)

[71.163 Run ID : 27097 - Run\_1-26\_12-11-29 276](#_Toc256000233)

[71.164 Run ID : 27096 - Run\_1-26\_12-10-13 276](#_Toc256000234)

[71.165 Run ID : 27095 - Fast\_Run\_1-26\_12-2-6 277](#_Toc256000235)

[71.166 Run ID : 27094 - Run\_1-26\_11-21-56 277](#_Toc256000236)

[71.167 Run ID : 27093 - Run\_1-26\_11-20-51 278](#_Toc256000237)

[71.168 Run ID : 27092 - Run\_1-26\_11-3-23 278](#_Toc256000238)

[71.169 Run ID : 27091 - Run\_1-26\_11-1-45 279](#_Toc256000239)

[71.170 Run ID : 27090 - Run\_1-26\_11-0-54 279](#_Toc256000240)

[71.171 Run ID : 27089 - Run\_1-26\_11-0-36 280](#_Toc256000241)

[71.172 Run ID : 27087 - Run\_1-25\_15-35-49 280](#_Toc256000242)

[71.173 Run ID : 27086 - Run\_1-25\_15-30-38 281](#_Toc256000243)

[71.174 Run ID : 27085 - Run\_1-25\_15-27-58 281](#_Toc256000244)

[71.175 Run ID : 27084 - Run\_1-25\_14-17-43 282](#_Toc256000245)

[71.176 Run ID : 27083 - Run\_1-25\_13-40-51 282](#_Toc256000246)

[71.177 Run ID : 27081 - Run\_1-25\_11-35-16 283](#_Toc256000247)

[71.178 Run ID : 27080 - Run\_1-25\_11-29-13 283](#_Toc256000248)

[71.179 Run ID : 27079 - Run\_1-25\_11-7-47 284](#_Toc256000249)

[71.180 Run ID : 27078 - Run\_1-25\_10-38-55 284](#_Toc256000250)

[71.181 Run ID : 27077 - Run\_1-25\_10-37-16 285](#_Toc256000251)

[71.182 Run ID : 27076 - Run\_1-24\_15-29-20 285](#_Toc256000252)

[71.183 Run ID : 27075 - Fast\_Run\_1-24\_15-29-20 285](#_Toc256000253)

[71.184 Run ID : 27073 - Run\_1-24\_11-46-10 286](#_Toc256000254)

[71.185 Run ID : 27071 - Run\_1-24\_11-35-30 286](#_Toc256000255)

[71.186 Run ID : 27069 - Run\_1-24\_9-9-45 287](#_Toc256000256)

[71.187 Run ID : 27068 - Run\_1-24\_9-7-37 287](#_Toc256000257)

[71.188 Run ID : 27067 - Run\_1-24\_9-6-54 288](#_Toc256000258)

[71.189 Run ID : 27027 - Run\_12-21\_10-54-33 288](#_Toc256000259)

[71.190 Run ID : 27022 - Fast\_Run\_12-21\_10-30-23 289](#_Toc256000260)

[71.191 Run ID : 27021 - Fast\_Run\_12-21\_10-30-20 289](#_Toc256000261)

[71.192 Run ID : 27020 - Fast\_Run\_12-21\_10-23-56 290](#_Toc256000262)

[71.193 Run ID : 27019 - Run\_12-21\_10-19-2 290](#_Toc256000263)

[71.194 Run ID : 27018 - Fast\_Run\_12-21\_10-18-35 291](#_Toc256000264)

[71.195 Run ID : 27017 - Run\_12-21\_9-53-42 291](#_Toc256000265)

[71.196 Run ID : 27013 - Fast\_Run\_12-21\_9-46-41 292](#_Toc256000266)

[71.197 Run ID : 27012 - Run\_12-21\_9-40-48 292](#_Toc256000267)

[71.198 Run ID : 27011 - Run\_12-20\_13-4-41 293](#_Toc256000268)

[71.199 Run ID : 27010 - Run\_12-20\_13-2-19 293](#_Toc256000269)

[71.200 Run ID : 27000 - Run\_12-19\_15-30-53 293](#_Toc256000270)

[71.201 Run ID : 26998 - Run\_12-19\_15-23-15 294](#_Toc256000271)

[71.202 Run ID : 26995 - Run\_12-19\_15-12-6 294](#_Toc256000272)

[71.203 Run ID : 26980 - Run\_12-19\_13-40-19 295](#_Toc256000273)

[71.204 Run ID : 26979 - Run\_12-19\_12-13-21 295](#_Toc256000274)

[71.205 Run ID : 26977 - Run\_12-19\_12-12-13 296](#_Toc256000275)

[71.206 Run ID : 26976 - Run\_12-19\_12-11-19 296](#_Toc256000276)

[71.207 Run ID : 26975 - Run\_12-19\_12-9-14 297](#_Toc256000277)

[71.208 Run ID : 26974 - Run\_12-19\_12-3-30 297](#_Toc256000278)

[71.209 Run ID : 26973 - Run\_12-19\_12-2-34 298](#_Toc256000279)

[71.210 Run ID : 26972 - Run\_12-19\_11-59-32 298](#_Toc256000280)

[71.211 Run ID : 26971 - Run\_12-19\_11-50-42 299](#_Toc256000281)

[71.212 Run ID : 26970 - Run\_12-19\_11-39-33 299](#_Toc256000282)

[71.213 Run ID : 26915 - Run\_12-15\_15-6-46 300](#_Toc256000283)

[71.214 Run ID : 26910 - Run\_12-15\_13-59-45 300](#_Toc256000284)

[71.215 Run ID : 26908 - Run\_12-15\_13-56-13 301](#_Toc256000285)

[71.216 Run ID : 26907 - Run\_12-15\_13-48-27 301](#_Toc256000286)

[71.217 Run ID : 26897 - Run\_12-15\_11-55-17 302](#_Toc256000287)

[71.218 Run ID : 26895 - Run\_12-15\_11-50-38 302](#_Toc256000288)

[71.219 Run ID : 26893 - Run\_12-15\_11-48-51 303](#_Toc256000289)

[71.220 Run ID : 26887 - Run\_12-15\_11-20-1 303](#_Toc256000290)

[71.221 Run ID : 26884 - Run\_12-15\_11-17-59 303](#_Toc256000291)

[71.222 Run ID : 26883 - Run\_12-15\_11-17-16 304](#_Toc256000292)

[71.223 Run ID : 26882 - Run\_12-15\_11-14-42 304](#_Toc256000293)

[71.224 Run ID : 26880 - Run\_12-15\_11-6-13 305](#_Toc256000294)

[71.225 Run ID : 26879 - Run\_12-15\_11-3-42 305](#_Toc256000295)

[71.226 Run ID : 26870 - Run\_12-15\_10-24-58 306](#_Toc256000296)

[71.227 Run ID : 26868 - Run\_12-15\_10-14-6 306](#_Toc256000297)

[71.228 Run ID : 26866 - Run\_12-15\_10-7-24 307](#_Toc256000298)

[71.229 Run ID : 26865 - Run\_12-15\_10-5-32 307](#_Toc256000299)

[71.230 Run ID : 26863 - Run\_12-15\_10-3-53 308](#_Toc256000300)

[71.231 Run ID : 26861 - Run\_12-15\_9-56-40 308](#_Toc256000301)

[71.232 Run ID : 26859 - Run\_12-15\_9-53-35 309](#_Toc256000302)

[71.233 Run ID : 26858 - Run\_12-15\_9-51-53 309](#_Toc256000303)

[71.234 Run ID : 26857 - Run\_12-15\_9-50-3 310](#_Toc256000304)

[71.235 Run ID : 26856 - Run\_12-15\_9-47-28 310](#_Toc256000305)

[71.236 Run ID : 26845 - Run\_12-14\_15-29-44 311](#_Toc256000306)

[71.237 Run ID : 26844 - Run\_12-14\_13-36-23 311](#_Toc256000307)

[71.238 Run ID : 26843 - Run\_12-14\_13-33-47 312](#_Toc256000308)

[71.239 Run ID : 26842 - Run\_12-14\_13-31-28 312](#_Toc256000309)

[71.240 Run ID : 26841 - Run\_12-14\_13-15-34 313](#_Toc256000310)

[71.241 Run ID : 26840 - Run\_12-14\_13-12-28 313](#_Toc256000311)

[71.242 Run ID : 26839 - Run\_12-14\_13-11-33 313](#_Toc256000312)

[71.243 Run ID : 26814 - Run\_12-12\_11-44-29 314](#_Toc256000313)

[71.244 Run ID : 26813 - Run\_12-12\_10-58-10 314](#_Toc256000314)

[71.245 Run ID : 26812 - Run\_12-8\_14-12-42 315](#_Toc256000315)

[71.246 Run ID : 26811 - Run\_12-8\_14-7-28 315](#_Toc256000316)

[71.247 Run ID : 26810 - Run\_12-8\_11-54-21 316](#_Toc256000317)

[71.248 Run ID : 26809 - Run\_12-8\_11-35-17 316](#_Toc256000318)

[71.249 Run ID : 26808 - Run\_12-8\_11-29-28 317](#_Toc256000319)

[71.250 Run ID : 26798 - Run\_10-28\_11-41-21 317](#_Toc256000320)

[71.251 Run ID : 26797 - Run\_10-28\_11-8-25 318](#_Toc256000321)

[71.252 Run ID : 26795 - Run\_10-28\_11-5-43 318](#_Toc256000322)

[71.253 Run ID : 26794 - Run\_10-28\_10-51-47 319](#_Toc256000323)

[71.254 Run ID : 26793 - Run\_10-28\_10-46-7 319](#_Toc256000324)

[71.255 Run ID : 26792 - Run\_10-28\_10-16-54 320](#_Toc256000325)

[71.256 Run ID : 26791 - Run\_10-28\_9-47-0 320](#_Toc256000326)

[71.257 Run ID : 26790 - Run\_10-27\_15-48-1 321](#_Toc256000327)

[71.258 Run ID : 26789 - Run\_10-27\_15-40-47 321](#_Toc256000328)

[71.259 Run ID : 26787 - Run\_10-27\_15-35-2 321](#_Toc256000329)

[71.260 Run ID : 26786 - Run\_10-27\_15-33-25 322](#_Toc256000330)

[71.261 Run ID : 26785 - Run\_10-27\_15-28-53 322](#_Toc256000331)

[71.262 Run ID : 26784 - Run\_10-27\_15-17-13 323](#_Toc256000332)

[71.263 Run ID : 26783 - Run\_10-27\_15-14-18 323](#_Toc256000333)

[71.264 Run ID : 26782 - Run\_10-27\_15-5-35 324](#_Toc256000334)

[71.265 Run ID : 26780 - Run\_10-27\_14-49-42 324](#_Toc256000335)

[71.266 Run ID : 26646 - Run\_10-17\_15-3-8 325](#_Toc256000336)

[71.267 Run ID : 26645 - Run\_10-17\_15-1-36 325](#_Toc256000337)

[71.268 Run ID : 26642 - Run\_10-17\_14-58-38 326](#_Toc256000338)

[71.269 Run ID : 26205 - Run\_9-29\_15-14-10 326](#_Toc256000339)

[71.270 Run ID : 26132 - Run\_9-23\_15-8-10 327](#_Toc256000340)

[71.271 Run ID : 26131 - Run\_9-23\_13-58-56 327](#_Toc256000341)

[71.272 Run ID : 26130 - Run\_9-23\_13-50-2 328](#_Toc256000342)

[71.273 Run ID : 26129 - Run\_9-23\_13-49-8 328](#_Toc256000343)

[71.274 Run ID : 26127 - Run\_9-23\_13-38-5 329](#_Toc256000344)

[71.275 Run ID : 26123 - Run\_9-23\_12-53-16 329](#_Toc256000345)

[71.276 Run ID : 26122 - Run\_9-23\_11-24-9 330](#_Toc256000346)

[71.277 Run ID : 26121 - Run\_9-23\_10-57-24 330](#_Toc256000347)

[71.278 Run ID : 26120 - Run\_9-23\_10-47-34 331](#_Toc256000348)

[71.279 Run ID : 26119 - Run\_9-23\_10-42-16 331](#_Toc256000349)

[71.280 Run ID : 26118 - Fast\_Run\_9-23\_10-42-13 332](#_Toc256000350)

[71.281 Run ID : 26117 - Run\_9-23\_10-33-31 332](#_Toc256000351)

[71.282 Run ID : 26116 - Run\_9-23\_10-30-25 333](#_Toc256000352)

[71.283 Run ID : 26115 - Run\_9-23\_10-0-13 333](#_Toc256000353)

[71.284 Run ID : 26114 - Run\_9-23\_9-57-13 334](#_Toc256000354)

[71.285 Run ID : 26113 - Run\_9-23\_9-29-12 334](#_Toc256000355)

[71.286 Run ID : 26112 - Run\_9-23\_9-25-24 335](#_Toc256000356)

[71.287 Run ID : 26111 - Run\_9-23\_9-7-54 335](#_Toc256000357)

[71.288 Run ID : 26108 - Run\_9-22\_14-18-12 336](#_Toc256000358)

[71.289 Run ID : 26096 - Run\_9-20\_14-41-8 336](#_Toc256000359)

# Homepage - General Screen

General testing of the MediPi homepage

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Call <MediPi Patient Banner (Called Test Only)>**

|  |
| --- |
| Description |
| Call <MediPi Patient Banner (Called Test Only)> |

|  |
| --- |
| Expected |
|  |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Confirm the configured tiles are displayed on the homepage |

|  |
| --- |
| Expected |
| The configured tiles are displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Tile Configuration

Testing MediPi tiles configuration i.e adding and removing tiles and checking functionality works as expected

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Make changes to the configuration to include only the finger Oximeter |

|  |
| --- |
| Expected |
| Tile displayed on homepage |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Take a reading and ensure it is displayed correctly on screen |

|  |
| --- |
| Expected |
| Reading displayed matches the device |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Navigate to the scheduler screen and ensure data can be submitted for the tiles available |

|  |
| --- |
| Expected |
| Scheduler allows data to be submitted for the tiles configured. |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| Navigate to the Clinical Web Front End |

|  |
| --- |
| Expected |
| Clinical web front end displayed with readings taken above |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Navigate to the transmitter screen and esnure data can be submitted for the tiles available |

|  |
| --- |
| Expected |
| Transmitter allows data to be submitted for the tiles configured |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 6**

|  |
| --- |
| Description |
| Navigate to the Clinical Web Front End |

|  |
| --- |
| Expected |
| Clinical web front end displayed with readings taken above |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 7**

|  |
| --- |
| Description |
| Navigate to the Notifications screen and ensure functions as expected |

|  |
| --- |
| Expected |
| Notifications screen functions as expected |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 9**

|  |
| --- |
| Description |
| Make changes to the configuration to add the Diagnostic Scale tile |

|  |
| --- |
| Expected |
| Tiles displayed on homepage |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 10**

|  |
| --- |
| Description |
| Repeat above steps with the additional tile added |

|  |
| --- |
| Expected |
| Results as above |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 11**

|  |
| --- |
| Description |
| Make changes to the configuration to add the Blood Pressure tile |

|  |
| --- |
| Expected |
| Tiles displayed on homepage |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 12**

|  |
| --- |
| Description |
| Repeat above steps with the additional tile added |

|  |
| --- |
| Expected |
| Results as above |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 13**

|  |
| --- |
| Description |
| Make changes to the configuration to add the Heart Questionnaire tile |

|  |
| --- |
| Expected |
| Tiles displayed on homepage |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 14**

|  |
| --- |
| Description |
| Repeat above steps with the additional tile added |

|  |
| --- |
| Expected |
| Results as above |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 15**

|  |
| --- |
| Description |
| Make changes to the configuration to add the Thermometer tile |

|  |
| --- |
| Expected |
| Tiles displayed on homepage |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 16**

|  |
| --- |
| Description |
| Repeat above steps with the additional tile added |

|  |
| --- |
| Expected |
| Results as above |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Logon - Admin Logon (Called Test Only)

General testing of the MediPi Logon screen

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Call <MediPi Patient Banner (Called Test Only)>**

|  |
| --- |
| Description |
| Call <MediPi Patient Banner (Called Test Only)> |

|  |
| --- |
| Expected |
|  |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Select the admin option |

|  |
| --- |
| Expected |
| Screen changes to allow for the entry of 10 digit password |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Select the cancel option without entering any characters |

|  |
| --- |
| Expected |
| Screen remains the same with the character fields blank |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| Reselect the admin option |

|  |
| --- |
| Expected |
| Screen changes to allow for a 10 digit password |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Populate nine characters of the password and select cancel |

|  |
| --- |
| Expected |
| The nine characters entered are cleared and the ten password fields are blank |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 6**

|  |
| --- |
| Description |
| Enter an incorrect password, once prevented select cancel |

|  |
| --- |
| Expected |
| User is denied access to MediPi and the password enrty fields go pink and only the cancel and close options are available |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 7**

|  |
| --- |
| Description |
| Enter a correct password via the logon screen, select Cancel to the prompt - Would you like to reboot into Configuration Mode? |

|  |
| --- |
| Expected |
| Screen changes to allow for a 10 digit password |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 8**

|  |
| --- |
| Description |
| Enter a correct password via the logon screen, select accept to the prompt - Would you like to reboot into Configuration Mode? |

|  |
| --- |
| Expected |
| User is granted access to the admin page of MediPi |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 9**

|  |
| --- |
| Description |
| Select the close option |

|  |
| --- |
| Expected |
| MediPi is closed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 10**

|  |
| --- |
| Description |
| Enter a correct password via the logon screen, select accept to the prompt - Would you like to reboot into Configuration Mode? |

|  |
| --- |
| Expected |
| User is granted access to the admin page of MediPi |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 11**

|  |
| --- |
| Description |
| Enter the password for the relevant user |

|  |
| --- |
| Expected |
| Access granted to MediPi Admin page with MediPi Settings tile only |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 12**

|  |
| --- |
| Description |
| Select the close option |

|  |
| --- |
| Expected |
| MediPi is closed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 10**

|  |
| --- |
| Description |
| Enter a correct password via the logon screen, select OK to the prompt - Would you like to reboot into Configuration Mode? |

|  |
| --- |
| Expected |
| User is granted access to the admin page of MediPi |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 11**

|  |
| --- |
| Description |
| Enter the password for the relevant user |

|  |
| --- |
| Expected |
| Access granted to MediPi Admin page with MediPi Settings tile only |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Logon - General and Patient Logon

General testing of the MediPi Logon screen

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Call <MediPi Patient Banner (Called Test Only)>**

|  |
| --- |
| Description |
| Call <MediPi Patient Banner (Called Test Only)> |

|  |
| --- |
| Expected |
|  |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Select the cancel option without entering any characters |

|  |
| --- |
| Expected |
| Screen remains the same with the character fields blank |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Select Admin option |

|  |
| --- |
| Expected |
| Ensure screen changes to allow for a 10 digit password |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| Select Patient option |

|  |
| --- |
| Expected |
| Ensure the screen changes back to allow for a 4 digit password |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Populate three characters of the password and select cancel |

|  |
| --- |
| Expected |
| The three characters entered are cleared and the four password fields are blank |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 6**

|  |
| --- |
| Description |
| Enter an incorrect password, once prevented select cancel |

|  |
| --- |
| Expected |
| User is denied access to MediPi and the password enrty fields go pink and only the cancel, Admin and close options are available |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 7**

|  |
| --- |
| Description |
| Enter a correct password via the logon screen |

|  |
| --- |
| Expected |
| User is granted access to MediPi and the homepage is displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 8**

|  |
| --- |
| Description |
| Select the close option |

|  |
| --- |
| Expected |
| MediPi is closed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Logon - MediPi Session Lock

Testing of the MediPi session lock

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Navigate to the following screens and allow MediPi to lock due to inactivity (300 Seconds) |

|  |
| --- |
| Expected |
| MediPi is locked and returns to the logon screen |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Finger Oximeter |

|  |
| --- |
| Expected |
|  |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Diagnostic Scale |

|  |
| --- |
| Expected |
|  |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| Blood Pressure |

|  |
| --- |
| Expected |
|  |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Heart Questionnaire |

|  |
| --- |
| Expected |
|  |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 6**

|  |
| --- |
| Description |
| Thermometer |

|  |
| --- |
| Expected |
|  |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 7**

|  |
| --- |
| Description |
| Scheduler |

|  |
| --- |
| Expected |
|  |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 8**

|  |
| --- |
| Description |
| Notifications |

|  |
| --- |
| Expected |
|  |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 9**

|  |
| --- |
| Description |
| Transmitter |

|  |
| --- |
| Expected |
|  |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 10**

|  |
| --- |
| Description |
| MediPi Settings via Admin Page |

|  |
| --- |
| Expected |
|  |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 11**

|  |
| --- |
| Description |
| Reconfigure the MediPi session timeout setting and repeat step 1 for each of the screens |

|  |
| --- |
| Expected |
| MediPi is locked and returns to the logon screen |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Admin Configuration Screen

PURPOSE: DEPENDENCIES/TEST DATA REF: PRE\_CONDITIONS:

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Call <Logon - Admin Logon (Called Test Only)>**

|  |
| --- |
| Description |
| Call <Logon - Admin Logon (Called Test Only)> |

|  |
| --- |
| Expected |
|  |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Select MediPi Settings |

|  |
| --- |
| Expected |
| MediPi admin screen displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Check the layout of the homepage is as per the approved design |

|  |
| --- |
| Expected |
| Layout as approved design |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| Enter more than 35 characters in the forename field and click save patient |

|  |
| --- |
| Expected |
| Error displayed - The forename/surname field must be 2 or more, but fewer than 35 characters long |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Select back option |

|  |
| --- |
| Expected |
| Ensure the user is returned to the admin homepage |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 6**

|  |
| --- |
| Description |
| Return to the admin page |

|  |
| --- |
| Expected |
| admin page displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 8**

|  |
| --- |
| Description |
| Select Bluetooth Serial Device dropdown |

|  |
| --- |
| Expected |
| Options displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 9**

|  |
| --- |
| Description |
| Enter less than 2 characters in the forename field and click save patient |

|  |
| --- |
| Expected |
| Error displayed - The forename/surname field must be 2 or more, but fewer than 35 characters long. |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 10**

|  |
| --- |
| Description |
| Enter 2 characters in the forename field and click save patient |

|  |
| --- |
| Expected |
| Forename saved successfully and displayed in the patient banner as entered |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 11**

|  |
| --- |
| Description |
| Enter 35 characters in the forename field and click save patient |

|  |
| --- |
| Expected |
| Forename saved successfully and displayed in the patient banner as entered |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 12**

|  |
| --- |
| Description |
| Enter more than 35 characters in the Surname field and click save patient |

|  |
| --- |
| Expected |
| Error displayed - The forename/surname field must be 2 or more, but fewer than 35 characters long |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 13**

|  |
| --- |
| Description |
| Enter less than 2 characters in the Surname field and click save patient |

|  |
| --- |
| Expected |
| Error displayed - The forename/surname field must be 2 or more, but fewer than 35 characters long. |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 14**

|  |
| --- |
| Description |
| Enter 2 characters in the Surname field and click save patient |

|  |
| --- |
| Expected |
| Surname saved successfully and displayed in the patient banner as entered |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 15**

|  |
| --- |
| Description |
| Enter 35 characters in the Surname field and click save patient |

|  |
| --- |
| Expected |
| Surname saved successfully and displayed in the patient banner as entered |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 16**

|  |
| --- |
| Description |
| Populate year field of date of birth and leave month and day fields blank |

|  |
| --- |
| Expected |
| Suitable error message displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 17**

|  |
| --- |
| Description |
| Populate month field of date of birth and leave year and day fields blank |

|  |
| --- |
| Expected |
| Suitable error message displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 18**

|  |
| --- |
| Description |
| Populate day field of date of birth and leave year and month fields blank |

|  |
| --- |
| Expected |
| Suitable error message displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 19**

|  |
| --- |
| Description |
| Enter a nine digit NHS Number and select Save Patient |

|  |
| --- |
| Expected |
| Suitable error message displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 20**

|  |
| --- |
| Description |
| Enter a 10 digit NHS Number but make it invalid and select Save Patient i.e. so it fails the checksum |

|  |
| --- |
| Expected |
| Suitable error message displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 21**

|  |
| --- |
| Description |
| Enter a valid NHS Number and select Save Patient |

|  |
| --- |
| Expected |
| NHS Number updated and displayed correctly in the patient banner |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 22**

|  |
| --- |
| Description |
| Enter an invalid MAC address in the Bluetooth Serial Port MAC field i.e. add an H to the end of the MAC address |

|  |
| --- |
| Expected |
| The Save BT Serial MAC button is disabled |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Finger Oximeter - General Screen

General testing of the Finger Oximeter screen

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Call <MediPi Patient Banner (Called Test Only)>**

|  |
| --- |
| Description |
| Call <MediPi Patient Banner (Called Test Only)> |

|  |
| --- |
| Expected |
|  |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Back and Start buttons displayed on screen as per approved design |

|  |
| --- |
| Expected |
| Back and Start button displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Back button returns user to the MedPi home page  Repeat from each of the instruction screens |

|  |
| --- |
| Expected |
| User returned to the MediPi home page  Results as above |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| MediPi displays - - BPM and - - SpO2 headings as per the approved design |

|  |
| --- |
| Expected |
| BPM and SpO2 headings displayed but with blank values until readings have been taken |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Navigate between device instructions by selecting left and right arrows |

|  |
| --- |
| Expected |
| Individual instructions displayed by selecting left and right arrows |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 6**

|  |
| --- |
| Description |
| 5 Separate instuctions displayed with relevant image |

|  |
| --- |
| Expected |
| Instructions and image displayed for each instruction |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Finger Oximeter - Record

Testing the record functionality from the finger oximeter and how this is displayed within MediPi

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Select Start without the device connected to MediPi    Repeat from each instruction screen |

|  |
| --- |
| Expected |
| User is requested to insert finger as per the instructions    Results as above |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Whilst taking a reading from the finger Oximeter remove finger from the device |

|  |
| --- |
| Expected |
| MediPi does not take any reading |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| Select start with the device connected and reading available    Confirm readings displayed on MediPi screen match those diisplayed via the Oximeter  Confirm the date and time measured is correct for the current date and time |

|  |
| --- |
| Expected |
| BPM and SP02 readings displayed on MediPi screen match those displayed on Oximeter  Date and time displayed correctly |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Navigate to homepage and check readings displayed against finger Oximeter tile match the last readings taken from the Oximeter |

|  |
| --- |
| Expected |
| Readings displayed against finger Oximeter tile match readings from Oximeter |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 6**

|  |
| --- |
| Description |
| Take a new reading via the finger Oximeter and ensure the updated readings are displayed within finger Oximeter MediPi screen |

|  |
| --- |
| Expected |
| Updated readings displayed within finger Oximeter MediPi screen |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 7**

|  |
| --- |
| Description |
| Navigate to homepage and check readings displayed against finger Oximeter tile match the last readings taken from the Oximeter |

|  |
| --- |
| Expected |
| Readings displayed against finger Oximeter tile match readings from Oximeter |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Finger Oximeter - Transmit

Testing transmitting data from MediPi to concentrator and how this is displayed on the clinical web application

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Take a reading via the Finger Oximeter |

|  |
| --- |
| Expected |
| reading displayed within MediPi and noted |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Navigate to Transmitter and select transmit to send the data to the concentrator |

|  |
| --- |
| Expected |
| Transmit selected and suitable confirmation displayed to the user |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Navigate to the Clinical Web Front End |

|  |
| --- |
| Expected |
| Clinical web front end displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| Check the BPM reading is displayed below the Pulse Rate heading |

|  |
| --- |
| Expected |
| BPM reading displayed below the Pulse Rate heading against the correct date and time |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Check the SpO2 reading is displayed below the Oxygen Saturation heading on the right handside as well as plotted on the chart |

|  |
| --- |
| Expected |
| Correct reading displayed below the Oxygen Saturation heading |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 6**

|  |
| --- |
| Description |
| Check as the medical devices transmit addiitonal data which is not displayed on the Clinical Web page has no impact on other data |

|  |
| --- |
| Expected |
| No impact on the Clinical Web page data or display |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Thermometer - General Screen

General testing of the Thermometer screen

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Call <MediPi Patient Banner (Called Test Only)>**

|  |
| --- |
| Description |
| Call <MediPi Patient Banner (Called Test Only)> |

|  |
| --- |
| Expected |
|  |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Back button displayed on screen as per approved design |

|  |
| --- |
| Expected |
| Back button displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Back button returns user to the MedPi home page    Repeat from each of the instruction screens |

|  |
| --- |
| Expected |
| User returned to the MediPi home page  Results as above |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| MediPi displays °c heading as per the approved design |

|  |
| --- |
| Expected |
| °c heading displayed but with blank value until reading has been taken |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Navigate between device instructions by selecting left and right arrows |

|  |
| --- |
| Expected |
| Individual instructions displayed by selecting left and right arrows |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 6**

|  |
| --- |
| Description |
| instructions displayed with relevant images |

|  |
| --- |
| Expected |
| Instruction and image displayed for instruction |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 7**

|  |
| --- |
| Description |
| Temperature can be entered by pressing any of the Plus or Minus buttons |

|  |
| --- |
| Expected |
| Temperature can be populated by pressing each of the buttons |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 8**

|  |
| --- |
| Description |
| three digit temperature can be entered by selecting relevant plus or minus options |

|  |
| --- |
| Expected |
| All Values can be populated |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 16**

|  |
| --- |
| Description |
| Confirm once all populated measured at time and date are populated |

|  |
| --- |
| Expected |
| Current time and date populated |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Thermometer - Record

Testing the record functionality from the Thermometer and how this is displayed within MediPi

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Navigate to thermometer screen |

|  |
| --- |
| Expected |
| Thermometer screen displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Leave reading as --.- and ensure nothing can be transmitted |

|  |
| --- |
| Expected |
| User is unable to transmit until a reading is provided |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Populate a valid reading and ensure correct measured at date and time is correct |

|  |
| --- |
| Expected |
| Reading populated with measured at with correct date and time |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Navigate to homepage and check readings displayed against thermometer tile match the last reading taken |

|  |
| --- |
| Expected |
| Readings displayed against thermometer tile match reading |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 6**

|  |
| --- |
| Description |
| Take a further reading and ensure correct measured at date and time is correct |

|  |
| --- |
| Expected |
| Reading populated with measured at with correct date and time |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 7**

|  |
| --- |
| Description |
| Navigate to homepage and check readings displayed against thermometer tile match the last reading taken |

|  |
| --- |
| Expected |
| Readings displayed against thermometer tile match reading |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Thermometer - Transmit

Testing transmitting data from MediPi to concentrator and how this is displayed on the clinical web application

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Take a reading via the thermometer screen |

|  |
| --- |
| Expected |
| reading displayed within MediPi and noted |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Navigate to Transmitter and select transmit to send the data to the concentrator |

|  |
| --- |
| Expected |
| Transmit selected and suitable confirmation displayed to the user |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Navigate to the Clinical Web Front End |

|  |
| --- |
| Expected |
| Clinical web front end displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| Check the °c reading is displayed below the Thermometer heading on the right handside as well as plotted on the chart |

|  |
| --- |
| Expected |
| °c reading displayed below the Thermometer heading against the correct date and time |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Check as the medical devices transmit addiitonal data which is not displayed on the Clinical Web page has no impact on other data |

|  |
| --- |
| Expected |
| No impact on the Clinical Web page data or display |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Diagnostic Scale - General Screen

General testing of the Diagnostic Scale screen

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Call <MediPi Patient Banner (Called Test Only)>**

|  |
| --- |
| Description |
| Call <MediPi Patient Banner (Called Test Only)> |

|  |
| --- |
| Expected |
|  |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Back and Start button displayed on screen as per approved design |

|  |
| --- |
| Expected |
| Back and Start button displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Back button returns user to the MedPi home page  Repeat from each of the instruction screens |

|  |
| --- |
| Expected |
| User returned to the MediPi home page  Results as above |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| MediPi displays - - Kg and Stone headings as per the approved design |

|  |
| --- |
| Expected |
| kg and st headings displayed but with blank values until readings has been taken |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Navigate between device instructions by selecting left and right arrows |

|  |
| --- |
| Expected |
| Individual instructions displayed by selecting left and right arrows |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 6**

|  |
| --- |
| Description |
| Separate instuctions displayed with relevant image |

|  |
| --- |
| Expected |
| Instructions and image displayed for each instruction |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Diagnostic Scale - Record

Testing the record functionality from the diagnostic scale and how this is displayed within MediPi

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Select Start without the device connected to MediPi    Repeat from each instruction screen |

|  |
| --- |
| Expected |
| User is asked to turn on the scales before proceeding  After a period of a time a warning is displayed to the user that the scales cannot be found and to check they are switched on and in range    Results as above |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Select Start with the device plugged in    Confirm readings displayed on MediPi screen match those taken from the scales |

|  |
| --- |
| Expected |
| Kg reading displayed on MediPi screen matches the reading taken from the scales |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| Navigate to homepage and check readings displayed against Diagnostic Scale tile match the last readings taken from the scales |

|  |
| --- |
| Expected |
| Readings displayed against Diagnostic Scale tile match readings from Scales |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Take a new reading via the scales and ensure the updated readings are displayed within Diagnostic Scale MediPi screen |

|  |
| --- |
| Expected |
| Updated readings displayed within Diagnostic Scale MediPi screen |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 6**

|  |
| --- |
| Description |
| Navigate to homepage and check readings displayed against Diagnostic Scale tile match the last readings taken from the Scales |

|  |
| --- |
| Expected |
| Readings displayed against Diagnostic Scale tile match readings from Scales |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Diagnostic Scale - Transmit

Testing transmitting data from MediPi to concentrator and how this is displayed on the clinical web application

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Take a reading via the Diagnostic Scale |

|  |
| --- |
| Expected |
| reading displayed within MediPi and noted |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Navigate to Transmitter and select transmit to send the data to the concentrator |

|  |
| --- |
| Expected |
| Transmit selected and suitable confirmation displayed to the user |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Navigate to the Clinical Web Front End |

|  |
| --- |
| Expected |
| Clinical web front end displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| Check the Kg reading is displayed below the Weight heading on the right handside as well as plotted on the chart |

|  |
| --- |
| Expected |
| Kg reading displayed below the Weight heading against the correct date and time |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 6**

|  |
| --- |
| Description |
| Check as the medical devices transmit addiitonal data which is not displayed on the Clinical Web page has no impact on other data |

|  |
| --- |
| Expected |
| No impact on the Clinical Web page data or display |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Blood Pressure - General Screen

General testing of the Blood Pressure screen

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Call <MediPi Patient Banner (Called Test Only)>**

|  |
| --- |
| Description |
| Call <MediPi Patient Banner (Called Test Only)> |

|  |
| --- |
| Expected |
|  |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Back and Start button displayed on screen as per approved design |

|  |
| --- |
| Expected |
| Back and Start button displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Back button returns user to the MedPi home page  Repeat from each of the instruction screens |

|  |
| --- |
| Expected |
| User returned to the MediPi home page  Results as above |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| MediPi displays - - mmHg, -- mmHg and BPM headings as per the approved design |

|  |
| --- |
| Expected |
| mmHg, mmHg and BPM headings displayed but with blank values until readings have been taken |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Navigate between device instructions by selecting left and right arrows |

|  |
| --- |
| Expected |
| Individual instructions displayed by selecting left and right arrows |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 6**

|  |
| --- |
| Description |
| Separate instuctions displayed with relevant images |

|  |
| --- |
| Expected |
| Instructions and images displayed for each instruction |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Blood Pressure - Record

Testing the record functionality from the Blood Pressure and how this is displayed within MediPi

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Select start without the device being connected to MediPi    Repeat from each instruction screen |

|  |
| --- |
| Expected |
| User is asked to press upload, before this happens the screen does not move    Results as above |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Upload a reading from the Blood Pressure device with not connection to Bluetooth |

|  |
| --- |
| Expected |
| Reading is not populated via MediPi |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Select start with the device plugged in    Confirm readings displayed on MediPi screen match those taken from the Blood Pressure Monitor |

|  |
| --- |
| Expected |
| mmHg, mmHg and BPM readings displayed on MediPi screen match the readings taken from the Blood Pressure Monitor |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| Navigate to homepage and check readings displayed against Blood Pressure tile match the last readings taken from the Blood Pressure Monitor |

|  |
| --- |
| Expected |
| Readings displayed against Blood Pressure tile match readings from Blood Pressure Monitor |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Take a new reading via the Blood Pressure Monitor and ensure the updated readings are displayed within Blood Pressure MediPi screen |

|  |
| --- |
| Expected |
| Updated readings displayed within Blood Pressure MediPi screen |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 6**

|  |
| --- |
| Description |
| Navigate to homepage and check readings displayed against Blood Pressure tile match the last readings taken from the Blood Pressure Monitor |

|  |
| --- |
| Expected |
| Readings displayed against Blood Pressure tile match readings from Blood Pressure Monitor |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Blood Pressure - Transmit

Testing transmitting data from MediPi to concentrator and how this is displayed on the clinical web application

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Take a reading via the Blood Pressure |

|  |
| --- |
| Expected |
| reading displayed within MediPi and noted |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Navigate to Transmitter and select transmit to send the data to the concentrator |

|  |
| --- |
| Expected |
| Transmit selected and suitable confirmation displayed to the user |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Navigate to the Clinical Web Front End |

|  |
| --- |
| Expected |
| Clinical web front end displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| Check the mmHg and mmHg readings are displayed below the Blood Pressure heading on the right handside as well as plotted on the chart |

|  |
| --- |
| Expected |
| mmHg and mmHg readings displayed below the Blood Pressure heading against the correct date and time |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 6**

|  |
| --- |
| Description |
| Check as the medical devices transmit addiitonal data which is not displayed on the Clinical Web page has no impact on other data |

|  |
| --- |
| Expected |
| No impact on the Clinical Web page data or display |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# COPD Questionnaire

General testing of the Questionnaire screen

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Call <MediPi Patient Banner (Called Test Only)>**

|  |
| --- |
| Description |
| Call <MediPi Patient Banner (Called Test Only)> |

|  |
| --- |
| Expected |
|  |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 1**

|  |
| --- |
| Description |
| Back and Start Questionnaire buttons displayed on screen as per approved design |

|  |
| --- |
| Expected |
| Back button and Start Questionnaire button displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Back button returns user to the MedPi home page |

|  |
| --- |
| Expected |
| User returned to the MediPi home page |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Select start questionanire to display first question, populate yes or now and select start questionnaire a second time |

|  |
| --- |
| Expected |
| The questionnaire is reset clearing previous results |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 6**

|  |
| --- |
| Description |
| Select start questionnaire and ensure first question is displayed  Repeat for remaining questions |

|  |
| --- |
| Expected |
| First question available with yes and no answers for the patient to populate  Results as per first question |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# COPD Questionnaire - Transmit

Testing transmitting data from MediPi to concentrator and how this is displayed on the clinical web application

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Populate answers to questions via the questionnaire |

|  |
| --- |
| Expected |
| Questions populated |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Navigate to Transmitter and select transmit to send the data to the concentrator |

|  |
| --- |
| Expected |
| Transmit selected and suitable confirmation displayed to the user |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# COPD Questionnaire Alerts

General testing of the Questionnaire screen

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Navigate to the COPD Questionnaire in MediPi |

|  |
| --- |
| Expected |
| COPD Questionnaire displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Using the attached spreadsheet navigate to the COPD Patient Interview tab |

|  |
| --- |
| Expected |
| COPD Patient Interview tab displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Test each of the Yes and No outcomes for each question ensuring the relevant alerts and questions are displayed based on the flow diagram |

|  |
| --- |
| Expected |
| Each question tested |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Diabetes Questionnaire

General testing of the Questionnaire screen

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Call <MediPi Patient Banner (Called Test Only)>**

|  |
| --- |
| Description |
| Call <MediPi Patient Banner (Called Test Only)> |

|  |
| --- |
| Expected |
|  |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 1**

|  |
| --- |
| Description |
| Back and Start Questionnaire buttons displayed on screen as per approved design |

|  |
| --- |
| Expected |
| Back button and Start Questionnaire button displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Back button returns user to the MedPi home page |

|  |
| --- |
| Expected |
| User returned to the MediPi home page |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Select start questionanire to display first question, populate yes or now and select start questionnaire a second time |

|  |
| --- |
| Expected |
| The questionnaire is reset clearing previous results |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 6**

|  |
| --- |
| Description |
| Select start questionnaire and ensure first question is displayed  Repeat for remaining questions |

|  |
| --- |
| Expected |
| First question available with yes and no answers for the patient to populate  Results as per first question |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Diabetes Questionnaire - Transmit

Testing transmitting data from MediPi to concentrator and how this is displayed on the clinical web application

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Populate answers to questions via the questionnaire |

|  |
| --- |
| Expected |
| Questions populated |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Navigate to Transmitter and select transmit to send the data to the concentrator |

|  |
| --- |
| Expected |
| Transmit selected and suitable confirmation displayed to the user |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Diabetes Questionnaire Alerts

General testing of the Questionnaire screen

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Navigate to the Diabetes Questionnaire in MediPi |

|  |
| --- |
| Expected |
| Diabetes Questionnaire displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Using the attached spreadsheet navigate to the Diabetes Patient Interview tab |

|  |
| --- |
| Expected |
| Diabetes Patient Interview tab displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Test each of the Yes and No outcomes for each question ensuring the relevant alerts and questions are displayed based on the flow diagram |

|  |
| --- |
| Expected |
| Each question tested |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Heart Failure Questionnaire

General testing of the Questionnaire screen

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Call <MediPi Patient Banner (Called Test Only)>**

|  |
| --- |
| Description |
| Call <MediPi Patient Banner (Called Test Only)> |

|  |
| --- |
| Expected |
|  |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 1**

|  |
| --- |
| Description |
| Back and Start Questionnaire buttons displayed on screen as per approved design |

|  |
| --- |
| Expected |
| Back button and Start Questionnaire button displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Back button returns user to the MedPi home page |

|  |
| --- |
| Expected |
| User returned to the MediPi home page |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Select start questionanire to display first question, populate yes or now and select start questionnaire a second time |

|  |
| --- |
| Expected |
| The questionnaire is reset clearing previous results |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 6**

|  |
| --- |
| Description |
| Select start questionnaire and ensure first question is displayed  Repeat for remaining questions |

|  |
| --- |
| Expected |
| First question available with yes and no answers for the patient to populate  Results as per first question |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Heart Questionnaire - Alerts

Testing of the Heart Questionnaire alerts based on results populated

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Navigate to the Heart Failure Questionnaire in MediPi |

|  |
| --- |
| Expected |
| Heart Failure Questionnaire displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Using the attached spreadsheet navigate to the Heart Failure Patient Interview tab |

|  |
| --- |
| Expected |
| Heart Failure Patient Interview tab displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Test each of the Yes and No outcomes for each question ensuring the relevant alerts and questions are displayed based on the flow diagram |

|  |
| --- |
| Expected |
| Each question tested |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Heart Questionnaire - Transmit

Testing transmitting data from MediPi to concentrator and how this is displayed on the clinical web application

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Populate answers to questions via the questionnaire |

|  |
| --- |
| Expected |
| Questions populated |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Navigate to Transmitter and select transmit to send the data to the concentrator |

|  |
| --- |
| Expected |
| Transmit selected and suitable confirmation displayed to the user |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Scheduler - Schedule Screens

General testing of the scheduler screen

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Call <MediPi Patient Banner (Called Test Only)>**

|  |
| --- |
| Description |
| Call <MediPi Patient Banner (Called Test Only)> |

|  |
| --- |
| Expected |
|  |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Back and Take Readings Now button displayed on screen as per approved design |

|  |
| --- |
| Expected |
| Back and Take Readings Now buttons displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Back button returns user to the MedPi home page |

|  |
| --- |
| Expected |
| User returned to the MediPi home page |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| Navigate to first schedule screen by selecting Take Readings Now    Confirm Back, Start and Next Options are available |

|  |
| --- |
| Expected |
| Patient navigated to finger Oximeter schedule screen  Options are available |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Call <Finger Oximeter - Record>**

|  |
| --- |
| Description |
| Call <Finger Oximeter - Record> |

|  |
| --- |
| Expected |
|  |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 6**

|  |
| --- |
| Description |
| Select Back |

|  |
| --- |
| Expected |
| Patient is returned to the MediPi homepage |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 7**

|  |
| --- |
| Description |
| Navigate back to finger Oximeter schedule screen and Select Next |

|  |
| --- |
| Expected |
| Ensure the patient is navigated to the next schedule screen |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 8**

|  |
| --- |
| Description |
| Confirm Back, start and Next Options are available |

|  |
| --- |
| Expected |
| Patient navigated to Diagnostic Scale schedule screen    Options are available |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Call <Diagnostic Scale - Record>**

|  |
| --- |
| Description |
| Call <Diagnostic Scale - Record> |

|  |
| --- |
| Expected |
|  |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 10**

|  |
| --- |
| Description |
| Select Cancel Schedule |

|  |
| --- |
| Expected |
| Patient is returned to the MediPi homepage |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 11**

|  |
| --- |
| Description |
| Navigate back to Diagnostic Scale schedule screen Select Next |

|  |
| --- |
| Expected |
| Ensure the patient is navigated to the next schedule screen |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 12**

|  |
| --- |
| Description |
| Confirm Back, Start and Next Options are available |

|  |
| --- |
| Expected |
| Patient navigated to Blood Pressure schedule screen    Options are available |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Call <Blood Pressure - Record>**

|  |
| --- |
| Description |
| Call <Blood Pressure - Record> |

|  |
| --- |
| Expected |
|  |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 14**

|  |
| --- |
| Description |
| Select Cancel Schedule |

|  |
| --- |
| Expected |
| Patient is returned to the MediPi homepage |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 15**

|  |
| --- |
| Description |
| Navigate back to Blood Pressure schedule screen Select Next |

|  |
| --- |
| Expected |
| Ensure the patient is navigated to the next schedule screen |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 16**

|  |
| --- |
| Description |
| Confirm Back and Next Options are available |

|  |
| --- |
| Expected |
| Patient navigated to Heart Questionnaire schedule screen    Options are available |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Call <Heart Failure Questionnaire>**

|  |
| --- |
| Description |
| Call <Heart Failure Questionnaire> |

|  |
| --- |
| Expected |
|  |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 18**

|  |
| --- |
| Description |
| Select Cancel Schedule |

|  |
| --- |
| Expected |
| Patient is returned to the MediPi homepage |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 19**

|  |
| --- |
| Description |
| Navigate back to Heart Questionnaire schedule screen Select Next |

|  |
| --- |
| Expected |
| Ensure the patient is navigated to the next schedule screen |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Call <Thermometer - Record>**

|  |
| --- |
| Description |
| Call <Thermometer - Record> |

|  |
| --- |
| Expected |
|  |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 21**

|  |
| --- |
| Description |
| Select Cancel Schedule |

|  |
| --- |
| Expected |
| Patient is returned to the MediPi homepage |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 22**

|  |
| --- |
| Description |
| Navigate back to Thermometer schedule screen Select Next |

|  |
| --- |
| Expected |
| Ensure the patient is navigated to the next schedule screen |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Scheduler - Transmit (All Data)

Testing transmitting all data from the scheduler screen to the concentrator and displayed on the clincial web application

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Navigate to the Scheduler |

|  |
| --- |
| Expected |
| Navigated to Scheduler |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Take a reading via the Finger Oximeter |

|  |
| --- |
| Expected |
| reading displayed within MediPi and noted |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Take a reading via the Diagnostic Scale |

|  |
| --- |
| Expected |
| reading displayed within MediPi and noted |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| Take a reading via the Blood Pressure |

|  |
| --- |
| Expected |
| reading displayed within MediPi and noted |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| Take a thermometer reading |

|  |
| --- |
| Expected |
| reading displayed within MediPi and noted |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Populate answers to questions via the questionnaire |

|  |
| --- |
| Expected |
| Questions populated |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 6**

|  |
| --- |
| Description |
| Select transmit to send the data to the concentrator |

|  |
| --- |
| Expected |
| Transmit selected and suitable confirmation displayed to the user |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 7**

|  |
| --- |
| Description |
| Navigate to the Clinical Web Front End |

|  |
| --- |
| Expected |
| Clinical web front end displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 8**

|  |
| --- |
| Description |
| Check the BPM reading is displayed below the Pulse Rate heading |

|  |
| --- |
| Expected |
| BPM reading displayed below the Pulse Rate heading against the correct date and time |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 9**

|  |
| --- |
| Description |
| Check the SpO2 reading is displayed below the Oxygen Saturation heading |

|  |
| --- |
| Expected |
| Correct reading displayed below the Oxygen Saturation heading against the correct date and time |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 10**

|  |
| --- |
| Description |
| Check the Kg reading is displayed below the Weight heading |

|  |
| --- |
| Expected |
| Kg reading displayed below the Weight heading against the correct date and time |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 11**

|  |
| --- |
| Description |
| Check the mmHg and mmHg readings are displayed below the Blood Pressure heading |

|  |
| --- |
| Expected |
| mmHg and mmHg readings displayed below the Blood Pressure heading against the correct date and time |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 11**

|  |
| --- |
| Description |
| Check the °c reading is displayed below the Thermometer heading |

|  |
| --- |
| Expected |
| °c reading displayed below the Thermometer heading against the correct date and time |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Scheduler - Transmit (Individual)

Testing transmitting individual data items from the transmitter screen to the concentrator and displayed on the clincial web application

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Navigate to the Scheduler |

|  |
| --- |
| Expected |
| Navigated to Scheduler |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Take a reading via the Finger Oximeter |

|  |
| --- |
| Expected |
| reading displayed within MediPi and noted |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Select transmit to send the data to the concentrator |

|  |
| --- |
| Expected |
| Transmit selected and suitable confirmation displayed to the user |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| Navigate to the Clinical Web Front End |

|  |
| --- |
| Expected |
| Clinical web front end displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Check the BPM reading is displayed below the Pulse Rate heading |

|  |
| --- |
| Expected |
| BPM reading displayed below the Pulse Rate heading against the correct date and time |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 6**

|  |
| --- |
| Description |
| Check the SpO2 reading is displayed below the Oxygen Saturation heading |

|  |
| --- |
| Expected |
| Correct reading displayed below the Oxygen Saturation heading against the correct date and time |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 7**

|  |
| --- |
| Description |
| Navigate back to the Scheduler |

|  |
| --- |
| Expected |
| Navigated to Scheduler |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 8**

|  |
| --- |
| Description |
| Take a reading via the Diagnostic Scale |

|  |
| --- |
| Expected |
| reading displayed within MediPi and noted |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 9**

|  |
| --- |
| Description |
| Select transmit to send the data to the concentrator |

|  |
| --- |
| Expected |
| Transmit selected and suitable confirmation displayed to the user |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 10**

|  |
| --- |
| Description |
| Navigate to the Clinical Web Front End |

|  |
| --- |
| Expected |
| Clinical web front end displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 11**

|  |
| --- |
| Description |
| Check the Kg reading is displayed below the Weight heading |

|  |
| --- |
| Expected |
| Kg reading displayed below the Weight heading against the correct date and time |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 12**

|  |
| --- |
| Description |
| Navigate back to the Scheduler |

|  |
| --- |
| Expected |
| Navigated to Scheduler |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 13**

|  |
| --- |
| Description |
| Take a reading via the Blood Pressure |

|  |
| --- |
| Expected |
| reading displayed within MediPi and noted |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 14**

|  |
| --- |
| Description |
| Select transmit to send the data to the concentrator |

|  |
| --- |
| Expected |
| Transmit selected and suitable confirmation displayed to the user |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 15**

|  |
| --- |
| Description |
| Navigate to the Clinical Web Front End |

|  |
| --- |
| Expected |
| Clinical web front end displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 16**

|  |
| --- |
| Description |
| Check the mmHg and mmHg readings are displayed below the Blood Pressure heading |

|  |
| --- |
| Expected |
| mmHg and mmHg readings displayed below the Blood Pressure heading against the correct date and time |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 17**

|  |
| --- |
| Description |
| Navigate back to the Scheduler |

|  |
| --- |
| Expected |
| Navigated to Scheduler |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 18**

|  |
| --- |
| Description |
| Populate answers to questions via the questionnaire |

|  |
| --- |
| Expected |
| Questions populated |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 19**

|  |
| --- |
| Description |
| Select transmit to send the data to the concentrator |

|  |
| --- |
| Expected |
| Transmit selected and suitable confirmation displayed to the user |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 20**

|  |
| --- |
| Description |
| Navigate back to the Scheduler |

|  |
| --- |
| Expected |
| Navigated to Scheduler |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 21**

|  |
| --- |
| Description |
| Take a thermometer reading |

|  |
| --- |
| Expected |
| reading displayed within MediPi and noted |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 22**

|  |
| --- |
| Description |
| Select transmit to send the data to the concentrator |

|  |
| --- |
| Expected |
| Transmit selected and suitable confirmation displayed to the user |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 23**

|  |
| --- |
| Description |
| Navigate to the Clinical Web Front End |

|  |
| --- |
| Expected |
| Clinical web front end displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 24**

|  |
| --- |
| Description |
| Check the °c reading is displayed below the Thermometer heading |

|  |
| --- |
| Expected |
| °c reading displayed below the Thermometer heading against the correct date and time |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Scheduler - Transmit Summary Screen

General testing of the transmit summary screen

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Confirm device data to be transmitted wording is displayed on screen as per approved design |

|  |
| --- |
| Expected |
| Wording displayed as expected |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Confirm the following tiles with associated tickboxes are displayed -    - Finger Oximeter  - Diagnostic Scale  - Blood Pressure (Sitting)  - Blood Pressure (Standing)  - Heart Questionnaire  - Diabetes Questionnaire  - COPD Questionnaire  - Thermometer  - Scheduler |

|  |
| --- |
| Expected |
|  |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Confirm Back and Transmit options are available |

|  |
| --- |
| Expected |
| Options are available |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| Select Back |

|  |
| --- |
| Expected |
| Patient is returned to the MediPi home page |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Navigate to the finger Oximeter schedule screen and take reading    Proceed to the Transmit Summary screen and ensure a tick is included against the finger Oximeter tile |

|  |
| --- |
| Expected |
| Navigated to finger Oximeter screen and reading taken  Transmit summary screen displayed with tickbox against the relevant tile |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 6**

|  |
| --- |
| Description |
| Repeat above step for remaining schedule screens i.e. diagnostic scale, blood pressure and heart questionnaire |

|  |
| --- |
| Expected |
| Results as per other screens |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 7**

|  |
| --- |
| Description |
| Select transmit without populating any of the scheduled screens |

|  |
| --- |
| Expected |
| MediPi prevents data being submitted until data has been populated |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 8**

|  |
| --- |
| Description |
| Take readings as part of a schedule and transmit the data to the server when its unavaiable |

|  |
| --- |
| Expected |
| Ensure the warning message displayed to the user is meaningful |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Scheduler Configuration

PURPOSE: DEPENDENCIES/TEST DATA REF: PRE\_CONDITIONS:

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Navigate to Scheduler and run the pre-configured schedule and transmit some readings to the Clinical Web Front end |

|  |
| --- |
| Expected |
| Reading required by the patient as part of pre-configured schedule i.e. once every day at 10am  Readings transmitted to the Clinical Web Application |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Via the MediPi configuration navigate to -  MediPi\_vx.x.xx                  Scheduler                                  Schedule.Schedule |

|  |
| --- |
| Expected |
| Navigated to relevant file in the MediPi configuration |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Navigate to the top of the file and amend column four on the first row identified with a status of Scheduled.    Edit the current value of 1440 to 60 |

|  |
| --- |
| Expected |
| File edited |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| Navigate to the bottom of the file and remove any rows which relate to today |

|  |
| --- |
| Expected |
| Lines removed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Log back onto MediPi and confirm that the scheduler has been reconfigured to require a reading based on the value set |

|  |
| --- |
| Expected |
| Reading required as per configuration setting |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 6**

|  |
| --- |
| Description |
| Navigate back to the MediPi configuration as above |

|  |
| --- |
| Expected |
| Navigated to relevant file in the MediPi configuration |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 7**

|  |
| --- |
| Description |
| Navigate to the top of the file and amend column four on the first row identified with a status of Scheduled.    Edit the current value of 60 to 1440 |

|  |
| --- |
| Expected |
| File edited |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Clinician Messages - General Screen

General testing of the clinician messages screen

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Call <MediPi Patient Banner (Called Test Only)>**

|  |
| --- |
| Description |
| Call <MediPi Patient Banner (Called Test Only)> |

|  |
| --- |
| Expected |
|  |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Back button displayed on screen as per approved design |

|  |
| --- |
| Expected |
| Back button displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Back button returns user to the MedPi home page |

|  |
| --- |
| Expected |
| User returned to the MediPi home page |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| Message list heading displayed |

|  |
| --- |
| Expected |
| Heading displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Message title and time displayed for each message |

|  |
| --- |
| Expected |
| Displayed as expected |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 6**

|  |
| --- |
| Description |
| Confirm the time is displayed in the format Fri 9 Sep 2016 13:56:48 BST |

|  |
| --- |
| Expected |
| Displayed in the correct format |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 7**

|  |
| --- |
| Description |
| Select one of the messages |

|  |
| --- |
| Expected |
| Message text displayed for the relevant message, i.e. the message with the correct title and date and time |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 8**

|  |
| --- |
| Description |
| Select another message and ensure the message text changes based on the message selected |

|  |
| --- |
| Expected |
| Relevant message text displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 9**

|  |
| --- |
| Description |
| Arrange for multiple messages to be loaded |

|  |
| --- |
| Expected |
| Ensure all messages are displayed correctly with relevant time/date message title and message text |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 10**

|  |
| --- |
| Description |
| Arrange for a message to be loaded with a really long message title and text, ensure the full message title and text are displayed without truncation |

|  |
| --- |
| Expected |
| Full message title and text displayed without truncation |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Clinician Messages - Polling Warning

Testing the polling of messages with connectivity unavailable

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Allow MediPi to poll the server for information whilst the server is unavailable |

|  |
| --- |
| Expected |
| Warning displayed is meaningful to the user |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Amend the configuration for the MediPi polling period |

|  |
| --- |
| Expected |
| Confirm MediPi polls based on the configuration value |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Transmitter - General Screen

General testing of the transmitter screen

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Call <MediPi Patient Banner (Called Test Only)>**

|  |
| --- |
| Description |
| Call <MediPi Patient Banner (Called Test Only)> |

|  |
| --- |
| Expected |
|  |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Back and transmit button displayed on screen as per approved design |

|  |
| --- |
| Expected |
| Back and transmit button displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Back button returns user to the MedPi home page |

|  |
| --- |
| Expected |
| User returned to the MediPi home page |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Confirm the following tiles with associated tickboxes are displayed -    - Finger Oximeter  - Diagnostic Scale  - Blood Pressure (Sitting)  - Blood Pressure (Standing)  - Thermometer  - Heart Questionnaire  - Diabetes Questionnaire  - COPD Questionnaire |

|  |
| --- |
| Expected |
| Tiles displayed with associated tickboxes |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 6**

|  |
| --- |
| Description |
| Confirm Transmit option is available |

|  |
| --- |
| Expected |
| Option is available |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 7**

|  |
| --- |
| Description |
| Navigate to the finger Oximeter screen and take reading    Proceed to the Transmit Summary screen and ensure a tick is included against the finger Oximeter tile |

|  |
| --- |
| Expected |
| Navigated to finger Oximeter screen and reading taken  Transmit summary screen displayed with tickbox against the relevant tile |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 8**

|  |
| --- |
| Description |
| Repeat above step for remaining screens i.e. diagnostic scale, blood pressure, heart questionnaire and thermometer |

|  |
| --- |
| Expected |
| Results as per other screens |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 9**

|  |
| --- |
| Description |
| Select transmit without populating any of the scheduled screens |

|  |
| --- |
| Expected |
| MediPi prevents data being submitted until data has been populated |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 10**

|  |
| --- |
| Description |
| Populate data and select transmit when the link to the server is unavailable |

|  |
| --- |
| Expected |
| Suitable warning message is displayed to the user |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 11**

|  |
| --- |
| Description |
| Populate data and select transmit when the link to the server is available |

|  |
| --- |
| Expected |
| Suitable confirmation message displayed to the user |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Transmitter - Transmit (All Data)

Testing transmitting all data from the transmitter screen to the concentrator and displayed on the clincial web application

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Take a reading via the Finger Oximeter |

|  |
| --- |
| Expected |
| reading displayed within MediPi and noted |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Take a reading via the Diagnostic Scale |

|  |
| --- |
| Expected |
| reading displayed within MediPi and noted |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Take a reading via the Blood Pressure |

|  |
| --- |
| Expected |
| reading displayed within MediPi and noted |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| Populate answers to questions via the questionnaires |

|  |
| --- |
| Expected |
| Questions populated |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Take a Thermometer reading |

|  |
| --- |
| Expected |
| reading displayed within MediPi and noted |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Navigate to Transmitter and select transmit to send the data to the concentrator |

|  |
| --- |
| Expected |
| Transmit selected and suitable confirmation displayed to the user |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 6**

|  |
| --- |
| Description |
| Navigate to the Clinical Web Front End |

|  |
| --- |
| Expected |
| Clinical web front end displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 7**

|  |
| --- |
| Description |
| Check the BPM reading is displayed below the Pulse Rate heading |

|  |
| --- |
| Expected |
| BPM reading displayed below the Pulse Rate heading against the correct date and time |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 8**

|  |
| --- |
| Description |
| Check the SpO2 reading is displayed below the Oxygen Saturation heading |

|  |
| --- |
| Expected |
| Correct reading displayed below the Oxygen Saturation heading against the correct date and time |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 9**

|  |
| --- |
| Description |
| Check the Kg reading is displayed below the Weight heading |

|  |
| --- |
| Expected |
| Kg reading displayed below the Weight heading against the correct date and time |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 10**

|  |
| --- |
| Description |
| Check the mmHg and mmHg readings are displayed below the Blood Pressure heading |

|  |
| --- |
| Expected |
| mmHg and mmHg readings displayed below the Blood Pressure heading against the correct date and time |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 10**

|  |
| --- |
| Description |
| Check the °c reading is displayed below the Thermometer heading |

|  |
| --- |
| Expected |
| °c reading displayed below the Thermometer heading against the correct date and time |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 11**

|  |
| --- |
| Description |
| Repeat above steps taking a further set of readings |

|  |
| --- |
| Expected |
| Additional readings recorded and transmitted |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 12**

|  |
| --- |
| Description |
| Navigate to the Clinical Web Front End |

|  |
| --- |
| Expected |
| Ensure additional readings are transmitted successfully from MediPi and displayed additionally against the previous reading against the correct date and time. |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 13**

|  |
| --- |
| Description |
| Repeat above steps taking a further set of readings the following day |

|  |
| --- |
| Expected |
| Additional readings recorded and transmitted the following day |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 14**

|  |
| --- |
| Description |
| Navigate to the Clinical Web Front End |

|  |
| --- |
| Expected |
| Ensure additional readings are transmitted successfully from MediPi and displayed additionally against the previous reading against the correct date and time. |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Logon & Logoff

PURPOSE: DEPENDENCIES/TEST DATA REF: PRE\_CONDITIONS:

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Attempt to access the Clinical Web Application by entering a username without a password |

|  |
| --- |
| Expected |
| Access is denied and suitable error message displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Attempt to access the Clinical Web Application by entering a password without a username |

|  |
| --- |
| Expected |
| Access is denied and suitable error message displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Attempt to access the Clinical Web Application by not entering a username or password |

|  |
| --- |
| Expected |
| Access is denied and suitable error message displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| Attempt to access the Clinical Web Application by entering a username with an invalid password |

|  |
| --- |
| Expected |
| Access is denied and suitable error message displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Attempt to access the Clinical Web Application by entering a username with a valid password |

|  |
| --- |
| Expected |
| Access granted successfully |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 6**

|  |
| --- |
| Description |
| Click the username top right of the screen and select logout |

|  |
| --- |
| Expected |
| The user is logged off the Clinical Web Application and must re-enter username and password to access again. |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 7**

|  |
| --- |
| Description |
| Re-run steps 1-5 to ensure functionality working as expected following log off |

|  |
| --- |
| Expected |
| Results as above |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Session Timeout

PURPOSE: DEPENDENCIES/TEST DATA REF: PRE\_CONDITIONS:

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Log onto the Clinical Web Application successfully |

|  |
| --- |
| Expected |
| Access granted |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Allow the web application to time out based on the time out settings |

|  |
| --- |
| Expected |
| Once the time out has been reached ensure the user's session is logged out |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Ensure to regain access, the user must enter their valid username and password |

|  |
| --- |
| Expected |
| Once entered access is granted |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Caseload View - Exclamation Mark

Error assessing any of the thresholds due to MediPi throwing an error. It is believed this can be tested by only populating one reading for weight in a given period.

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Log onto MediPi and generate readings/data for    - Oximeter  - Scales  - Blood Pressure  - Thermometer  - Heart Questionnnaire    Do not transmit |

|  |
| --- |
| Expected |
| Readings/data populated and noted |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Navigate to the Clinical Web Page via Google Chrome, select Manage > Patients links and select relevant patient tile |

|  |
| --- |
| Expected |
| Navigated to the patient detail view for the relevant patient |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Set the following threshold values for    - Oximeter - within range  - Scales - period set to 72 hours with weight change set to 2  - Blood Pressure - within range  - Thermometer - within range  - Heart Questionnnaire |

|  |
| --- |
| Expected |
| Thresholds set |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| Navigate to MediPi and transmit data via transmitter |

|  |
| --- |
| Expected |
| Data transmitted via transmitter |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Navigate to the caseload view page and ensure for the relevant patient the exclamation mark icon is displayed as only one weight reading has been recorded |

|  |
| --- |
| Expected |
| Exclamation mark displayed as all values are in range |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 6**

|  |
| --- |
| Description |
| Select the relevant patient tile and ensure readings populated above are displayed in green |

|  |
| --- |
| Expected |
| Readings are correct and displayed in green |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 7**

|  |
| --- |
| Description |
| Repeat above steps by transmitting data via the scheduler rather than transmitter |

|  |
| --- |
| Expected |
| Results as above |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Caseload View - Happy Face

Everything submitted in time and and in range

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Log onto MediPi and generate readings/data for    - Oximeter  - Scales  - Blood Pressure  - Thermometer  - Heart Questionnnaire    Do not transmit |

|  |
| --- |
| Expected |
| Readings/data populated and noted |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Navigate to the Clinical Web Page via Google Chrome, select Manage > Patients links and select relevant patient tile |

|  |
| --- |
| Expected |
| Navigated to the patient detail view for the relevant patient |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Set the threshold values for    - Oximeter  - Scales  - Blood Pressure  - Thermometer  - Heart Questionnnaire    Set accordingly to ensure the readings taken are in range |

|  |
| --- |
| Expected |
| Thresholds set |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| Navigate to MediPi and transmit data via transmitter |

|  |
| --- |
| Expected |
| Data transmitted via transmitter |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Navigate to the caseload view page and ensure for the relevant patient the happy face icon is displayed as all values are in range |

|  |
| --- |
| Expected |
| Happy face displayed as all values are in range |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 6**

|  |
| --- |
| Description |
| Select the relevant patient tile and ensure readings populated above are displayed in green |

|  |
| --- |
| Expected |
| Readings are correct and displayed in green |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 7**

|  |
| --- |
| Description |
| Repeat above steps by transmitting data via the scheduler rather than transmitter |

|  |
| --- |
| Expected |
| Results as above |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Caseload View - Question Mark

Everything submitted in time and something out of range

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Log onto MediPi and ensure no readings/data have been transmitted for the given period |

|  |
| --- |
| Expected |
| No readings taken for the given period |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Navigate to the Clinical Web Page via Google Chrome, select Manage > Patients links and select relevant patient tile |

|  |
| --- |
| Expected |
| Navigated to the patient detail view for the relevant patient |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Navigate to the caseload view page and ensure for the relevant patient the question mark icon is displayed as no readings have been recorded for the given period |

|  |
| --- |
| Expected |
| Question mark displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Caseload View - Question Mark Reset

Error assessing any of the thresholds due to MediPi throwing an error. It is believed this can be tested by only populating one reading for weight in a given period.

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Log onto MediPi and generate readings/data for    - Oximeter  - Scales  - Blood Pressure  - Thermometer  - Heart Questionnnaire    Do not transmit |

|  |
| --- |
| Expected |
| Readings/data populated and noted |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Navigate to the Clinical Web Page via Google Chrome, select Manage > Patients links and select relevant patient tile |

|  |
| --- |
| Expected |
| Navigated to the patient detail view for the relevant patient |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Set the threshold values for    - Oximeter  - Scales  - Blood Pressure  - Thermometer  - Heart Questionnnaire    Set accordingly to ensure the readings taken are in range |

|  |
| --- |
| Expected |
| Thresholds set |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| Navigate to MediPi and transmit data via transmitter |

|  |
| --- |
| Expected |
| Data transmitted via transmitter |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Navigate to the caseload view page and ensure for the relevant patient the happy face icon is displayed as all values are in range |

|  |
| --- |
| Expected |
| Happy face displayed as all values are in range |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 6**

|  |
| --- |
| Description |
| Select the relevant patient tile and ensure readings populated above are displayed in green |

|  |
| --- |
| Expected |
| Readings are correct and displayed in green |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 7**

|  |
| --- |
| Description |
| Either let the scheduled event period reset or amend the value via scheduler configuration file |

|  |
| --- |
| Expected |
| Scheduled event period reset |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 8**

|  |
| --- |
| Description |
| Navigate to the caseload view page and ensure for the relevant patient the happy face icon has now changed to a question mark as no readings have been recorded for the given period |

|  |
| --- |
| Expected |
| Happy face changed to a question mark |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Caseload View - Sad Face

Everything submitted in time and and in range

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Log onto MediPi and generate readings/data for    - Oximeter  - Scales  - Blood Pressure  - Thermometer  - Heart Questionnnaire    Do not transmit |

|  |
| --- |
| Expected |
| Readings/data populated and noted |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Navigate to the Clinical Web Page via Google Chrome, select Manage > Patients links and select relevant patient tile |

|  |
| --- |
| Expected |
| Navigated to the patient detail view for the relevant patient |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Set the threshold values for    - Oximeter  - Scales  - Blood Pressure  - Thermometer  - Heart Questionnnaire    Set accordingly to ensure the readings taken are out of range |

|  |
| --- |
| Expected |
| Thresholds set |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| Navigate to MediPi and transmit data via transmitter |

|  |
| --- |
| Expected |
| Data transmitted via transmitter |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Navigate to the caseload view page and ensure for the relevant patient the sad face icon is displayed as all values are out of range |

|  |
| --- |
| Expected |
| Sad face displayed as all values are out range |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 6**

|  |
| --- |
| Description |
| Select the relevant patient tile and ensure readings populated above are displayed in red |

|  |
| --- |
| Expected |
| Readings are correct and displayed in red |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 7**

|  |
| --- |
| Description |
| Repeat above steps by amending the thresholds to send some data within range and some out of range |

|  |
| --- |
| Expected |
| Results as above |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 8**

|  |
| --- |
| Description |
| Repeat above steps by transmitting data via the scheduler rather than transmitter |

|  |
| --- |
| Expected |
| Results as above |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Browser Testing - Internet Explorer

Clinical Web Application tested in Internet Explorer

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Repeat all Clinical Web Application tests using Internet Explorer |

|  |
| --- |
| Expected |
| All tests repeated |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Homepage - General Screen

General testing of the Clinical Web Application homepage

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Navigate to the Clinical Web Page via Google Chrome |

|  |
| --- |
| Expected |
| Clinical Web Page displayed in Google Chrome |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Check the layout of the homepage is as per the approved design |

|  |
| --- |
| Expected |
| Layout as approved design |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Select all links from the homepage |

|  |
| --- |
| Expected |
| All links selected |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Patient Detail View

General testing of the patient detail view screen

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Navigate to the Clinical Web Page via Google Chrome |

|  |
| --- |
| Expected |
| Clinical Web Page displayed in Google Chrome |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Select one of the patient tiles |

|  |
| --- |
| Expected |
| Navigated to the patient details view |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Check the layout of the homepage is as per the approved design |

|  |
| --- |
| Expected |
| Layout as approved design |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| Check the patient banner displayed is as per the approved design |

|  |
| --- |
| Expected |
| Patient banner displayed as per approved design |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Check the patient details displayed match the tile selected from the caseload view screen |

|  |
| --- |
| Expected |
| Details match |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 6**

|  |
| --- |
| Description |
| Check that the patient banner is displayed at all times at the top of the screen |

|  |
| --- |
| Expected |
| Banner displayed at all times |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 7**

|  |
| --- |
| Description |
| Select all links |

|  |
| --- |
| Expected |
| All links selected |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 8**

|  |
| --- |
| Description |
| Minimise each of the sections |

|  |
| --- |
| Expected |
| Each section on screen is collapsed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 9**

|  |
| --- |
| Description |
| Expand each of the sections |

|  |
| --- |
| Expected |
| Each section expanded |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 10**

|  |
| --- |
| Description |
| Remove and add each of the plotted lines against each section |

|  |
| --- |
| Expected |
| Minimum, maximum and data lines can be added and removed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 11**

|  |
| --- |
| Description |
| Repeat above steps for another patient tile on the caseload view screen |

|  |
| --- |
| Expected |
| Results as above |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Patients Caseload View

General testing of the patient's caseload view screen

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Navigate to the Clinical Web Page via Google Chrome |

|  |
| --- |
| Expected |
| Clinical Web Page displayed in Google Chrome |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 1**

|  |
| --- |
| Description |
| Select Manage > Patients links |

|  |
| --- |
| Expected |
| Navigated to the caseload view |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Check the layout of the homepage is as per the approved design |

|  |
| --- |
| Expected |
| Layout as approved design |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Select all links |

|  |
| --- |
| Expected |
| All links selected |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| Check for each patient a tile is displayed and layout is as per approved Design |

|  |
| --- |
| Expected |
| Each patient has a separate tile and layout correct |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Select the first patient tile and check the patient details on the caseload view match those on the patient detail |

|  |
| --- |
| Expected |
| User navigated to the patient detail for the patient selected and details match |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 6**

|  |
| --- |
| Description |
| Select back on the browser |

|  |
| --- |
| Expected |
| User navigated back to the caseload view |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 7**

|  |
| --- |
| Description |
| Select the second patient tile and check the patient details on the caseload view match those on the patient detail |

|  |
| --- |
| Expected |
| User navigated to the patient detail for the patient selected and details match |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 8**

|  |
| --- |
| Description |
| Select Patients link from the menu |

|  |
| --- |
| Expected |
| User navigated back to the caseload view |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 9**

|  |
| --- |
| Description |
| Repeat for remaining tiles on screen |

|  |
| --- |
| Expected |
| User navigated to the patient details for each of the tiles selected |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Data transmitted from multiple MediPi Devices

PURPOSE: DEPENDENCIES/TEST DATA REF: PRE\_CONDITIONS:

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Logon to two MediPi Devices |

|  |
| --- |
| Expected |
| Two MediPi Devices logged on |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Ensure a set of medical devices are connected to each of the MediPi devices |

|  |
| --- |
| Expected |
| Medical devices connected |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Take a set of readings using the first MediPi device  Note the readings against the patient configured on the device |

|  |
| --- |
| Expected |
| Readings and patient details recorded |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| Take a set of readings using the second MediPi device  Note the readings against the patient configured on the device |

|  |
| --- |
| Expected |
| Readings and patient details recorded |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Log on to the Clinical Web Application |

|  |
| --- |
| Expected |
| Clinical Web Application displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 6**

|  |
| --- |
| Description |
| Transmit data from each of the MediPi devices |

|  |
| --- |
| Expected |
| Data transmitted |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 7**

|  |
| --- |
| Description |
| Via the Clinical Web Application check the data transmitted from the first MediPi device is displayed as expected against the correct patient |

|  |
| --- |
| Expected |
| Data displayed correctly for the relevant patient |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 8**

|  |
| --- |
| Description |
| Via the Clinical Web Application check the data transmitted from the second MediPi device is displayed as expected against the correct patient |

|  |
| --- |
| Expected |
| Data displayed correctly for the relevant patient |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Fatal Error Screen

Testing of the fatal error screen

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Arrange for a fatal error to be displayed i.e. amend the configuration file |

|  |
| --- |
| Expected |
| Check the warning displayed to the user is meaningful |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Check that once the user is on the fatal error screen that can navigate back to the MediPi homepage |

|  |
| --- |
| Expected |
| User is able to navigate back to MediPi homepage |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Loss of WIFI Connection

PURPOSE: DEPENDENCIES/TEST DATA REF: PRE\_CONDITIONS:

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Access the Raspberry Pi desktop by either typing startx from the command prompt or by logging onto the admin section of MediPi and selecting the top toolbar. |

|  |
| --- |
| Expected |
| Raspberry Pi desktop displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Turn off the WIFI so the Raspberry PI has no connection to the internet |

|  |
| --- |
| Expected |
| WIFI switched off on the Raspberry Pi |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Reboot the Raspberry Pi and logon to MediPi |

|  |
| --- |
| Expected |
| Raspberry rebooted and the MediPi logon screen displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| Confirm at the bottom of the screen that a message is displayed - WIFI Connection lost    Other warnings are likely to display due to the network being unreachable  Data cannot be transmitted to the concentrator |

|  |
| --- |
| Expected |
| Warning messages displayed and data cannot be transmitted to the concentrator |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Turn on the WIFI so the Raspberry PI has a connection to the internet |

|  |
| --- |
| Expected |
| WIFI switched on |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 6**

|  |
| --- |
| Description |
| Confirm the WIFI Connection lost message is no longer displayed at the bottom of the screen |

|  |
| --- |
| Expected |
| Message no longer displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 7**

|  |
| --- |
| Description |
| Transmit a couple of readings from MediPi |

|  |
| --- |
| Expected |
| Confirm data can be transmitted to the concentrator |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# MediPi Patient Banner - Patient Name

Testing of a long patient name displayed on the MediPi Patient Banner

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Log onto MediPi with a patient surname of 35 characters |

|  |
| --- |
| Expected |
| MediPi displays Full patient surname displayed without truncation. General screen layout has not impacted by the long name |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Log onto MediPi with a patient forename of 35 characters |

|  |
| --- |
| Expected |
| Full patient forename displayed without truncation General screen layout has not impacted by the long name |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# MediPi Patient Banner (Called Test Only)

General testing of the MediPi Patient Banner

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

|  |
| --- |
| Description |
| MedPi Title displayed as per approved design |

|  |
| --- |
| Expected |
| MediPi title displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

|  |
| --- |
| Description |
| Patients Surname and Forename displayed as per approved design |

|  |
| --- |
| Expected |
| Patient's Surname and Forename displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

|  |
| --- |
| Description |
| Patient's Date of Birth displayed as per approved design |

|  |
| --- |
| Expected |
| Date of Birth displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

|  |
| --- |
| Description |
| Patient's NHS Number displayed as per approved design |

|  |
| --- |
| Expected |
| NHS Number displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

|  |
| --- |
| Description |
| Time and Date displayed as per approved design |

|  |
| --- |
| Expected |
| Date and Time displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

|  |
| --- |
| Description |
| Connection icon displayed as per approved design |

|  |
| --- |
| Expected |
| Connection icon displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

|  |
| --- |
| Description |
| Confirm the red cross is displayed top right of the screen |

|  |
| --- |
| Expected |
| Red cross displayed top right |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

|  |
| --- |
| Description |
| Select the red cross and ensure MediPi is closed |

|  |
| --- |
| Expected |
| MediPi is closed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Multiple Connected Devices

Testing MediPi with multiple connected devices

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Connect all bluetooth devices to MediPi |

|  |
| --- |
| Expected |
| Ensure each of the devices can be used and readings recorded to MediPi |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# MediPi Application - Order of readings changed via schedule

PURPOSE: DEPENDENCIES/TEST DATA REF: PRE\_CONDITIONS:

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Logon to MediPi Application |

|  |
| --- |
| Expected |
| Logged on successfully |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Navigate to the readings tile |

|  |
| --- |
| Expected |
| Reading screens displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Navigate through the different readings and confirm the order is as follows -    - BP Sitting  - Pulse  - Temperature  - BP Standing  - Weight |

|  |
| --- |
| Expected |
| Order of the readings correct |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# MediPi Application - Removal of all tiles on the home screen except for the readings tile

PURPOSE: DEPENDENCIES/TEST DATA REF: PRE\_CONDITIONS:

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Logon to the MediPi application |

|  |
| --- |
| Expected |
| MediPi homepage displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Confirm the only tile displayed is the readings tile |

|  |
| --- |
| Expected |
| Only one tile displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Confirm the readings tile can be accessed and readings successfully transmitted to the MediPi Clinical Web Application |

|  |
| --- |
| Expected |
| Readings recorded and transmitted to the MediPi Clinical Web Application |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| Confirm all readings via the MediPi Clinical Web Application are showing in either red or green |

|  |
| --- |
| Expected |
| Readings show in red or green |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Return to the MediPi Clinical Web Application tomorrow and confirm the tiles on the patient screen are displayed in grey due to no readings being taken for the given period |

|  |
| --- |
| Expected |
| Readings are showing in grey until scheduled readings from the MediPi application have been transmitted |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# MediPi Application - Removal of PID on main MediPi Screen before authentication

PURPOSE: DEPENDENCIES/TEST DATA REF: PRE\_CONDITIONS:

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Navigate to the MediPi application logon screen |

|  |
| --- |
| Expected |
| Username and password fields displayed on MediPi application homepage |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Confirm that the patient forename and surname is displayed in the banner but not the NHS Number and date of birth |

|  |
| --- |
| Expected |
| Patient forename and surname displayed on the patient banner    NHS Number and date of birth are not displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Enter a valid username and password to access the application |

|  |
| --- |
| Expected |
| Access granted to MediPi application |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| Confirm thay the patient forename, surname, NHS Number and date of birth are displayed in the patient banner |

|  |
| --- |
| Expected |
| Patient forename, surname, NHS Number and date of birth are displayed on patient banner following authentication |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Log off the MediPi application and confirm only the forename and surname are now displayed in the patient banner |

|  |
| --- |
| Expected |
| Patient forename and surname displayed on the patient banner    NHS Number and date of birth are not displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# MediPi Application - Timeout increased to 15 minutes

PURPOSE: DEPENDENCIES/TEST DATA REF: PRE\_CONDITIONS:

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Log onto the MediPi application |

|  |
| --- |
| Expected |
| Logged onto the MediPi application successfully |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Leave the application for 15 mins |

|  |
| --- |
| Expected |
| After 15 mins the application should timeout and the user returned to the logon screen |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Users must re-authenticate before accessing the application after the session has timed out |

|  |
| --- |
| Expected |
| User unable to access the application before a valid username and password have been entered |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# MediPi Clinical Application - Weight thresholds changed to be simple high & low

PURPOSE: DEPENDENCIES/TEST DATA REF: PRE\_CONDITIONS:

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Log on to the MediPi Clinical Web Application and the MediPi application |

|  |
| --- |
| Expected |
| Logged onto to MediPi Clinical Web Application and MediPi Application |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Take a weight reading via the scales and transmit the reading to the MediPi Clinical Web Application |

|  |
| --- |
| Expected |
| Reading taken via the scales and transmitted to the MediPi Clinical Web Application  Note the reading |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Navigate to the weight reading via the MediPi Clinical Web Application and amend the low value to be above the reading taken above |

|  |
| --- |
| Expected |
| Weight readings displayed, weight low threshold amended to be above the reading above |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| Take a second weight reading via the scales and transmit the reading to the MediPi Clinical Web Application |

|  |
| --- |
| Expected |
| Reading taken and transmitted |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Return to the MediPi Clinical Web Application and confirm the reading transmitted is showing in red due to be out of range of the low threshold |

|  |
| --- |
| Expected |
| Reading showing in red |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 6**

|  |
| --- |
| Description |
| Navigate to the weight reading via the MediPi Clinical Web Application and amend the high value to be below the reading taken above |

|  |
| --- |
| Expected |
| Weight readings displayed, weight low threshold amended to be below the reading above |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 7**

|  |
| --- |
| Description |
| Take a second weight reading via the scales and transmit the reading to the MediPi Clinical Web Application |

|  |
| --- |
| Expected |
| Reading taken and transmitted |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 8**

|  |
| --- |
| Description |
| Return to the MediPi Clinical Web Application and confirm the reading transmitted is showing in red due to be out of range of the high threshold |

|  |
| --- |
| Expected |
| Reading showing in red |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 9**

|  |
| --- |
| Description |
| Navigate to the weight reading via the MediPi Clinical Web Application and set the high and low threshold values to be in range with the reading taken above |

|  |
| --- |
| Expected |
| Weight readings displayed, weight low and high threshold amended to be below the reading above |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 10**

|  |
| --- |
| Description |
| Take a second weight reading via the scales and transmit the reading to the MediPi Clinical Web Application |

|  |
| --- |
| Expected |
| Reading taken and transmitted |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 11**

|  |
| --- |
| Description |
| Return to the MediPi Clinical Web Application and confirm the reading transmitted is showing in green as the reading is now within range |

|  |
| --- |
| Expected |
| Reading showing in green |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Blood Pressure Threshold

Testing of the blood pressure thresholds i.e. out of range, within range

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Log onto MediPi and generate a reading but don't transmit |

|  |
| --- |
| Expected |
| Reading populated and noted |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Navigate to the Clinical Web Page via Google Chrome, select Manage > Patients links and select relevant patient tile |

|  |
| --- |
| Expected |
| Navigated to the patient detail view for the relevant patient |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Amend the minimum threshold value to above the reading taken above |

|  |
| --- |
| Expected |
| Minimum threshold value amended |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| Navigate back to MediPi and transmit the reading taken above |

|  |
| --- |
| Expected |
| Ensure an alert is generated to MediPi due to the reading being out of range. |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Navigate back to the Clinical Web Page, ensure the reading above has been successfully transmitted, plotted on screen below the minimum line and the correct reading is displayed in red |

|  |
| --- |
| Expected |
| Reading transmitted successfully, plotted on screen and reading displayed in red |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 6**

|  |
| --- |
| Description |
| Navigate back to MediPi and generate a reading but don't transmit |

|  |
| --- |
| Expected |
| Reading populated and noted |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 7**

|  |
| --- |
| Description |
| Navigate to the Clinical Web Page via Google Chrome, select Manage > Patients links and select relevant patient tile |

|  |
| --- |
| Expected |
| Navigated to the patient detail view for the relevant patient |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 8**

|  |
| --- |
| Description |
| Amend the maximum threshold value to above the reading taken above |

|  |
| --- |
| Expected |
| Maximum threshold value amended |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 9**

|  |
| --- |
| Description |
| Navigate back to MediPi and transmit the reading taken above |

|  |
| --- |
| Expected |
| Ensure an alert is generated to MediPi due to the reading being out of range. |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 10**

|  |
| --- |
| Description |
| Navigate back to the Clinical Web Page, ensure the reading above has been successfully transmitted, plotted on screen above the maximum line and the correct reading is displayed in red |

|  |
| --- |
| Expected |
| Reading transmitted successfully, plotted on screen and reading displayed in red |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 11**

|  |
| --- |
| Description |
| Navigate back to MediPi and generate a reading but don't transmit |

|  |
| --- |
| Expected |
| Reading populated and noted |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 12**

|  |
| --- |
| Description |
| Navigate to the Clinical Web Page via Google Chrome, select Manage > Patients links and select relevant patient tile |

|  |
| --- |
| Expected |
| Navigated to the patient detail view for the relevant patient |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 13**

|  |
| --- |
| Description |
| Amend the minimum and maximum threshold values to ensure the reading taken above is in range |

|  |
| --- |
| Expected |
| Minimum and Maximum threshold values amended |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 14**

|  |
| --- |
| Description |
| Navigate back to MediPi and transmit the reading taken above |

|  |
| --- |
| Expected |
| Ensure an alert is not generated due to the reading being within range |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 15**

|  |
| --- |
| Description |
| Navigate back to the Clinical Web Page, ensure the reading above has been successfully transmitted, plotted on screen between the minimum and maximum line and the correct reading is displayed in green |

|  |
| --- |
| Expected |
| Reading transmitted successfully, plotted on screen and reading displayed in green |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 16**

|  |
| --- |
| Description |
| Ensure the minimum and maximum lines are displayed with the correct values |

|  |
| --- |
| Expected |
| Minimum and Maximum lines displayed with correct values |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 17**

|  |
| --- |
| Description |
| Repeat steps where the minimum threshold is set to the value provided by the device |

|  |
| --- |
| Expected |
| Ensure the reading is identified as within range |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 18**

|  |
| --- |
| Description |
| Repeat steps where the maximum threshold is set to the value provided by the device |

|  |
| --- |
| Expected |
| Ensure the reading is identified as within range |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# COPD Questionnaire Results

PURPOSE: DEPENDENCIES/TEST DATA REF: PRE\_CONDITIONS:

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Populate answers to questions via the questionnaire to generate a positive result |

|  |
| --- |
| Expected |
| Questions populated |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Navigate to Transmitter and select transmit to send the data to the concentrator |

|  |
| --- |
| Expected |
| Transmit selected and suitable confirmation displayed to the user |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Log onto the Clinical Web Application and ensure a green column with the current date/time has been added to the COPD Questionnaire section |

|  |
| --- |
| Expected |
| Green column added |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| Hover over the created green column and check to ensure the questions are populated as generated from MediPi |

|  |
| --- |
| Expected |
| Questions populated with correct outcomes  Date and time included above the questions/answers |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Repeat above steps but generate a negative outcome, ensure a red column is created |

|  |
| --- |
| Expected |
| Red column created with the correct outcomes against the questions |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Diabetes Questionnaire Results

PURPOSE: DEPENDENCIES/TEST DATA REF: PRE\_CONDITIONS:

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Populate answers to questions via the questionnaire to generate a positive result |

|  |
| --- |
| Expected |
| Questions populated |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Navigate to Transmitter and select transmit to send the data to the concentrator |

|  |
| --- |
| Expected |
| Transmit selected and suitable confirmation displayed to the user |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Log onto the Clinical Web Application and ensure a green column with the current date/time has been added to the Diabetes Questionnaire section |

|  |
| --- |
| Expected |
| Green column added |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| Hover over the created green column and check to ensure the questions are populated as generated from MediPi |

|  |
| --- |
| Expected |
| Questions populated with correct outcomes  Date and time included above the questions/answers |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Repeat above steps but generate a negative outcome, ensure a red column is created |

|  |
| --- |
| Expected |
| Red column created with the correct outcomes against the questions |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Heart Questionnaire Results

PURPOSE: DEPENDENCIES/TEST DATA REF: PRE\_CONDITIONS:

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Populate answers to questions via the questionnaire to generate a positive result |

|  |
| --- |
| Expected |
| Questions populated |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Navigate to Transmitter and select transmit to send the data to the concentrator |

|  |
| --- |
| Expected |
| Transmit selected and suitable confirmation displayed to the user |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Log onto the Clinical Web Application and ensure a green column with the current date/time has been added to the Heart Questionnaire section |

|  |
| --- |
| Expected |
| Green column added |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| Hover over the created green column and check to ensure the questions are populated as generated from MediPi |

|  |
| --- |
| Expected |
| Questions populated with correct outcomes  Date and time included above the questions/answers |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Repeat above steps but generate a negative outcome, ensure a red column is created |

|  |
| --- |
| Expected |
| Red column created with the correct outcomes against the questions |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Oxygen Saturation Threshold

Testing of the Oxygen Saturation thresholds i.e. out of range, within range

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Log onto MediPi and generate a reading using Oximeter but don't transmit |

|  |
| --- |
| Expected |
| Reading populated and noted |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Navigate to the Clinical Web Page via Google Chrome, select Manage > Patients links and select relevant patient tile |

|  |
| --- |
| Expected |
| Navigated to the patient detail view for the relevant patient |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Amend the minimum threshold value to above the reading taken above |

|  |
| --- |
| Expected |
| Minimum threshold value amended |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| Navigate back to MediPi and transmit the reading taken above |

|  |
| --- |
| Expected |
| Ensure an alert is generated to MediPi due to the reading being out of range. |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Navigate back to the Clinical Web Page, ensure the reading above has been successfully transmitted, plotted on screen below the minimum line and the correct reading is displayed in red |

|  |
| --- |
| Expected |
| Reading transmitted successfully, plotted on screen and reading displayed in red |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 6**

|  |
| --- |
| Description |
| Navigate back to MediPi and generate a reading via Oximeter but don't transmit |

|  |
| --- |
| Expected |
| Reading populated and noted |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 7**

|  |
| --- |
| Description |
| Navigate to the Clinical Web Page via Google Chrome, select Manage > Patients links and select relevant patient tile |

|  |
| --- |
| Expected |
| Navigated to the patient detail view for the relevant patient |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 8**

|  |
| --- |
| Description |
| Amend the maximum threshold value to above the reading taken above |

|  |
| --- |
| Expected |
| Maximum threshold value amended |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 9**

|  |
| --- |
| Description |
| Navigate back to MediPi and transmit the reading taken above |

|  |
| --- |
| Expected |
| Ensure an alert is generated to MediPi due to the reading being out of range. |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 10**

|  |
| --- |
| Description |
| Navigate back to the Clinical Web Page, ensure the reading above has been successfully transmitted, plotted on screen above the maximum line and the correct reading is displayed in red |

|  |
| --- |
| Expected |
| Reading transmitted successfully, plotted on screen and reading displayed in red |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 11**

|  |
| --- |
| Description |
| Navigate back to MediPi and generate a reading via Oximeter but don't transmit |

|  |
| --- |
| Expected |
| Reading populated and noted |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 12**

|  |
| --- |
| Description |
| Navigate to the Clinical Web Page via Google Chrome, select Manage > Patients links and select relevant patient tile |

|  |
| --- |
| Expected |
| Navigated to the patient detail view for the relevant patient |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 13**

|  |
| --- |
| Description |
| Amend the minimum and maximum threshold values to ensure the reading taken above is in range |

|  |
| --- |
| Expected |
| Minimum and Maximum threshold values amended |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 14**

|  |
| --- |
| Description |
| Navigate back to MediPi and transmit the reading taken above |

|  |
| --- |
| Expected |
| Ensure an alert is generated due to the reading being within range |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 15**

|  |
| --- |
| Description |
| Navigate back to the Clinical Web Page, ensure the reading above has been successfully transmitted, plotted on screen between the minimum and maximum line and the correct reading is displayed in green |

|  |
| --- |
| Expected |
| Reading transmitted successfully, plotted on screen and reading displayed in green |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 16**

|  |
| --- |
| Description |
| Ensure the minimum and maximum lines are displayed with the correct values |

|  |
| --- |
| Expected |
| Minimum and Maximum lines displayed with correct values |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 17**

|  |
| --- |
| Description |
| Repeat steps where the minimum threshold is set to the value provided by the device |

|  |
| --- |
| Expected |
| Ensure the reading is identified as within range |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 18**

|  |
| --- |
| Description |
| Repeat steps where the maximum threshold is set to the value provided by the device |

|  |
| --- |
| Expected |
| Ensure the reading is identified as within range |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Pulse Rate Threshold

Testing of the Pulse Rate thresholds i.e. out of range, within range

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Log onto MediPi and generate a reading but don't transmit |

|  |
| --- |
| Expected |
| Reading populated and noted |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Navigate to the Clinical Web Page via Google Chrome, select Manage > Patients links and select relevant patient tile |

|  |
| --- |
| Expected |
| Navigated to the patient detail view for the relevant patient |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Amend the minimum threshold value to above the reading taken above |

|  |
| --- |
| Expected |
| Minimum threshold value amended |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| Navigate back to MediPi and transmit the reading taken above |

|  |
| --- |
| Expected |
| Ensure an alert is generated to MediPi due to the reading being out of range. |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Navigate back to the Clinical Web Page, ensure the reading above has been successfully transmitted, plotted on screen below the minimum line and the correct reading is displayed in red |

|  |
| --- |
| Expected |
| Reading transmitted successfully, plotted on screen and reading displayed in red |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 6**

|  |
| --- |
| Description |
| Navigate back to MediPi and generate a reading but don't transmit |

|  |
| --- |
| Expected |
| Reading populated and noted |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 7**

|  |
| --- |
| Description |
| Navigate to the Clinical Web Page via Google Chrome, select Manage > Patients links and select relevant patient tile |

|  |
| --- |
| Expected |
| Navigated to the patient detail view for the relevant patient |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 8**

|  |
| --- |
| Description |
| Amend the maximum threshold value to below the reading taken above |

|  |
| --- |
| Expected |
| Maximum threshold value amended |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 9**

|  |
| --- |
| Description |
| Navigate back to MediPi and transmit the reading taken above |

|  |
| --- |
| Expected |
| Ensure an alert is generated to MediPi due to the reading being out of range. |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 10**

|  |
| --- |
| Description |
| Navigate back to the Clinical Web Page, ensure the reading above has been successfully transmitted, plotted on screen above the maximum line and the correct reading is displayed in red |

|  |
| --- |
| Expected |
| Reading transmitted successfully, plotted on screen and reading displayed in red |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 11**

|  |
| --- |
| Description |
| Navigate back to MediPi and generate a reading but don't transmit |

|  |
| --- |
| Expected |
| Reading populated and noted |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 12**

|  |
| --- |
| Description |
| Navigate to the Clinical Web Page via Google Chrome, select Manage > Patients links and select relevant patient tile |

|  |
| --- |
| Expected |
| Navigated to the patient detail view for the relevant patient |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 13**

|  |
| --- |
| Description |
| Amend the minimum and maximum threshold values to ensure the reading taken above is in range |

|  |
| --- |
| Expected |
| Minimum and Maximum threshold values amended |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 14**

|  |
| --- |
| Description |
| Navigate back to MediPi and transmit the reading taken above |

|  |
| --- |
| Expected |
| Ensure an alert is not generated due to the reading being within range |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 15**

|  |
| --- |
| Description |
| Navigate back to the Clinical Web Page, ensure the reading above has been successfully transmitted, plotted on screen between the minimum and maximum line and the correct reading is displayed in green |

|  |
| --- |
| Expected |
| Reading transmitted successfully, plotted on screen and reading displayed in green |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 16**

|  |
| --- |
| Description |
| Ensure the minimum and maximum lines are displayed with the correct values |

|  |
| --- |
| Expected |
| Minimum and Maximum lines displayed with correct values |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 17**

|  |
| --- |
| Description |
| Repeat steps where the minimum threshold is set to the value provided by the device |

|  |
| --- |
| Expected |
| Ensure the reading is identified as within range |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 18**

|  |
| --- |
| Description |
| Repeat steps where the maximum threshold is set to the value provided by the device |

|  |
| --- |
| Expected |
| Ensure the reading is identified as within range |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Temperature Threshold

Testing of the Temperature thresholds i.e. out of range, within range

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Log onto MediPi and generate a reading but don't transmit |

|  |
| --- |
| Expected |
| Reading populated and noted |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Navigate to the Clinical Web Page via Google Chrome, select Manage > Patients links and select relevant patient tile |

|  |
| --- |
| Expected |
| Navigated to the patient detail view for the relevant patient |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Amend the minimum threshold value to above the reading taken above |

|  |
| --- |
| Expected |
| Minimum threshold value amended |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| Navigate back to MediPi and transmit the reading taken above |

|  |
| --- |
| Expected |
| Ensure an alert is generated to MediPi due to the reading being out of range. |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Navigate back to the Clinical Web Page, ensure the reading above has been successfully transmitted, plotted on screen below the minimum line and the correct reading is displayed in red |

|  |
| --- |
| Expected |
| Reading transmitted successfully, plotted on screen and reading displayed in red |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 6**

|  |
| --- |
| Description |
| Navigate back to MediPi and generate a reading but don't transmit |

|  |
| --- |
| Expected |
| Reading populated and noted |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 7**

|  |
| --- |
| Description |
| Navigate to the Clinical Web Page via Google Chrome, select Manage > Patients links and select relevant patient tile |

|  |
| --- |
| Expected |
| Navigated to the patient detail view for the relevant patient |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 8**

|  |
| --- |
| Description |
| Amend the maximum threshold value to below the reading taken above |

|  |
| --- |
| Expected |
| Maximum threshold value amended |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 9**

|  |
| --- |
| Description |
| Navigate back to MediPi and transmit the reading taken above |

|  |
| --- |
| Expected |
| Ensure an alert is generated to MediPi due to the reading being out of range. |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 10**

|  |
| --- |
| Description |
| Navigate back to the Clinical Web Page, ensure the reading above has been successfully transmitted, plotted on screen above the maximum line and the correct reading is displayed in red |

|  |
| --- |
| Expected |
| Reading transmitted successfully, plotted on screen and reading displayed in red |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 11**

|  |
| --- |
| Description |
| Navigate back to MediPi and generate a reading but don't transmit |

|  |
| --- |
| Expected |
| Reading populated and noted |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 12**

|  |
| --- |
| Description |
| Navigate to the Clinical Web Page via Google Chrome, select Manage > Patients links and select relevant patient tile |

|  |
| --- |
| Expected |
| Navigated to the patient detail view for the relevant patient |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 13**

|  |
| --- |
| Description |
| Amend the minimum and maximum threshold values to ensure the reading taken above is in range |

|  |
| --- |
| Expected |
| Minimum and Maximum threshold values amended |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 14**

|  |
| --- |
| Description |
| Navigate back to MediPi and transmit the reading taken above |

|  |
| --- |
| Expected |
| Ensure an alert is not generated due to the reading being within range |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 15**

|  |
| --- |
| Description |
| Navigate back to the Clinical Web Page, ensure the reading above has been successfully transmitted, plotted on screen between the minimum and maximum line and the correct reading is displayed in green |

|  |
| --- |
| Expected |
| Reading transmitted successfully, plotted on screen and reading displayed in green |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 16**

|  |
| --- |
| Description |
| Ensure the minimum and maximum lines are displayed with the correct values |

|  |
| --- |
| Expected |
| Minimum and Maximum lines displayed with correct values |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 17**

|  |
| --- |
| Description |
| Repeat steps where the minimum threshold is set to the value provided by the device |

|  |
| --- |
| Expected |
| Reading identified within range and no alerts generated |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 18**

|  |
| --- |
| Description |
| Repeat steps where the maximum threshold is set to the value provided by the device |

|  |
| --- |
| Expected |
| Reading identified within range and no alerts generated |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Threshold Input Value Validation

Testing the threshold input value validation to ensure values outside the range cannot be entered

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Select Manage > Patients links |

|  |
| --- |
| Expected |
| Navigated to the caseload view |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Select a patient tile |

|  |
| --- |
| Expected |
| Navigated to the patient detail view for the relevant patient |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Temperature  Attempt to enter a lower value than is supported in the minimum value field |

|  |
| --- |
| Expected |
| User is unable to enter |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| Attempt to enter a higher value than is supported in the minimum value field |

|  |
| --- |
| Expected |
| User is unable to enter |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Attempt to enter a lower value than is supported in the maximum value field |

|  |
| --- |
| Expected |
| User is unable to enter |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 6**

|  |
| --- |
| Description |
| Attempt to enter a higher value than is supported in the maximum value field |

|  |
| --- |
| Expected |
| User is unable to enter |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 7**

|  |
| --- |
| Description |
| Enter a valid minimum and maximum value |

|  |
| --- |
| Expected |
| Ensure the threshold values have been updated successfully |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 8**

|  |
| --- |
| Description |
| Pulse Rate  Attempt to enter a lower value than is supported in the minimum value field |

|  |
| --- |
| Expected |
| User is unable to enter |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 9**

|  |
| --- |
| Description |
| Attempt to enter a higher value than is supported in the minimum value field |

|  |
| --- |
| Expected |
| User is unable to enter |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 10**

|  |
| --- |
| Description |
| Attempt to enter a lower value than is supported in the maximum value field |

|  |
| --- |
| Expected |
| User is unable to enter |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 11**

|  |
| --- |
| Description |
| Attempt to enter a higher value than is supported in the maximum value field |

|  |
| --- |
| Expected |
| User is unable to enter |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 12**

|  |
| --- |
| Description |
| Enter a valid minimum and maximum value |

|  |
| --- |
| Expected |
| Ensure the threshold values have been updated successfully |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 13**

|  |
| --- |
| Description |
| Weight  Attempt to enter a lower value than is supported in the minimum value field |

|  |
| --- |
| Expected |
| User is unable to enter |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 14**

|  |
| --- |
| Description |
| Attempt to enter a higher value than is supported in the minimum value field |

|  |
| --- |
| Expected |
| User is unable to enter |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 15**

|  |
| --- |
| Description |
| Attempt to enter a lower value than is supported in the maximum value field |

|  |
| --- |
| Expected |
| User is unable to enter |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 16**

|  |
| --- |
| Description |
| Attempt to enter a higher value than is supported in the maximum value field |

|  |
| --- |
| Expected |
| User is unable to enter |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 17**

|  |
| --- |
| Description |
| Enter a valid minimum and maximum value |

|  |
| --- |
| Expected |
| Ensure the threshold values have been updated successfully |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 18**

|  |
| --- |
| Description |
| Oxygen Saturation  Attempt to enter a lower value than is supported in the minimum value field |

|  |
| --- |
| Expected |
| User is unable to enter |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 19**

|  |
| --- |
| Description |
| Attempt to enter a higher value than is supported in the minimum value field |

|  |
| --- |
| Expected |
| User is unable to enter |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 20**

|  |
| --- |
| Description |
| Attempt to enter a lower value than is supported in the maximum value field |

|  |
| --- |
| Expected |
| User is unable to enter |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 21**

|  |
| --- |
| Description |
| Attempt to enter a higher value than is supported in the maximum value field |

|  |
| --- |
| Expected |
| User is unable to enter |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 22**

|  |
| --- |
| Description |
| Enter a valid minimum and maximum value |

|  |
| --- |
| Expected |
| Ensure the threshold values have been updated successfully |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 23**

|  |
| --- |
| Description |
| Blood Pressure  Attempt to enter a lower value than is supported in the minimum value field |

|  |
| --- |
| Expected |
| User is unable to enter |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 24**

|  |
| --- |
| Description |
| Attempt to enter a higher value than is supported in the minimum value field |

|  |
| --- |
| Expected |
| User is unable to enter |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 25**

|  |
| --- |
| Description |
| Attempt to enter a lower value than is supported in the maximum value field |

|  |
| --- |
| Expected |
| User is unable to enter |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 26**

|  |
| --- |
| Description |
| Attempt to enter a higher value than is supported in the maximum value field |

|  |
| --- |
| Expected |
| User is unable to enter |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 27**

|  |
| --- |
| Description |
| Enter a valid minimum and maximum value |

|  |
| --- |
| Expected |
| Ensure the threshold values have been updated successfully |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 28**

|  |
| --- |
| Description |
| Repeat for an addiitonal patient |

|  |
| --- |
| Expected |
| Results as above |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Weight Threshold

Testing of the weight thresholds i.e. out of range, within range

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Log onto MediPi and generate a reading but don't transmit |

|  |
| --- |
| Expected |
| Reading populated and noted |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Navigate to the Clinical Web Page via Google Chrome, select Manage > Patients links and select relevant patient tile |

|  |
| --- |
| Expected |
| Navigated to the patient detail view for the relevant patient |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Amend the minimum threshold value to above the reading taken above |

|  |
| --- |
| Expected |
| Minimum threshold value amended |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| Navigate back to MediPi and transmit the reading taken above |

|  |
| --- |
| Expected |
| Ensure an alert is generated to MediPi due to the reading being out of range. |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Navigate back to the Clinical Web Page, ensure the reading above has been successfully transmitted, plotted on screen below the minimum line and the correct reading is displayed in red |

|  |
| --- |
| Expected |
| Reading transmitted successfully, plotted on screen and reading displayed in red |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 6**

|  |
| --- |
| Description |
| Navigate back to MediPi and generate a reading but don't transmit |

|  |
| --- |
| Expected |
| Reading populated and noted |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 7**

|  |
| --- |
| Description |
| Navigate to the Clinical Web Page via Google Chrome, select Manage > Patients links and select relevant patient tile |

|  |
| --- |
| Expected |
| Navigated to the patient detail view for the relevant patient |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 8**

|  |
| --- |
| Description |
| Amend the maximum threshold value to above the reading taken above |

|  |
| --- |
| Expected |
| Maximum threshold value amended |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 9**

|  |
| --- |
| Description |
| Navigate back to MediPi and transmit the reading taken above |

|  |
| --- |
| Expected |
| Ensure an alert is generated to MediPi due to the reading being out of range. |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 10**

|  |
| --- |
| Description |
| Navigate back to the Clinical Web Page, ensure the reading above has been successfully transmitted, plotted on screen above the maximum line and the correct reading is displayed in red |

|  |
| --- |
| Expected |
| Reading transmitted successfully, plotted on screen and reading displayed in red |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 11**

|  |
| --- |
| Description |
| Navigate back to MediPi and generate a reading but don't transmit |

|  |
| --- |
| Expected |
| Reading populated and noted |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 12**

|  |
| --- |
| Description |
| Navigate to the Clinical Web Page via Google Chrome, select Manage > Patients links and select relevant patient tile |

|  |
| --- |
| Expected |
| Navigated to the patient detail view for the relevant patient |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 13**

|  |
| --- |
| Description |
| Amend the minimum and maximum threshold values to ensure the reading taken above is in range |

|  |
| --- |
| Expected |
| Minimum and Maximum threshold values amended |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 14**

|  |
| --- |
| Description |
| Navigate back to MediPi and transmit the reading taken above |

|  |
| --- |
| Expected |
| Ensure an alert is not generated due to the reading being within range |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 15**

|  |
| --- |
| Description |
| Navigate back to the Clinical Web Page, ensure the reading above has been successfully transmitted, plotted on screen between the minimum and maximum line and the correct reading is displayed in green |

|  |
| --- |
| Expected |
| Reading transmitted successfully, plotted on screen and reading displayed in green |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 16**

|  |
| --- |
| Description |
| Ensure the minimum and maximum lines are displayed with the correct values |

|  |
| --- |
| Expected |
| Minimum and Maximum lines displayed with correct values |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 17**

|  |
| --- |
| Description |
| Repeat steps where the minimum threshold is set to the value provided by the device |

|  |
| --- |
| Expected |
| Ensure the reading is identified as within range |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 18**

|  |
| --- |
| Description |
| Repeat steps where the maximum threshold is set to the value provided by the device |

|  |
| --- |
| Expected |
| Ensure the reading is identified as within range |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Blood Pressure Monitor - Time Synchronisation

PURPOSE: DEPENDENCIES/TEST DATA REF: PRE\_CONDITIONS:

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Reset the time on the Blood pressure monitor    This can be achieved by removing the batteries for 30 seconds or amending the time within the settings |

|  |
| --- |
| Expected |
| Time reset to incorrect on the finger oximeter |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Take a reading via the blood pressure monitor and a screen is displayed to provide instructions for reseting the time via the blood pressure monitor  Reset the date and time to current time |

|  |
| --- |
| Expected |
| Time and date against the blood pressure monitor reset |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Retake the reading via the blood pressure monitor and upload the reading to MediPi    Check the reading and the date and time is correct |

|  |
| --- |
| Expected |
| Reading taken displayed correctly in MediPi against the current date and time |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| Transmit the reading to the clinical web application |

|  |
| --- |
| Expected |
| Reading transmitted to the clinical web application and displayed against the correct date and time |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Finger Oximeter - Time Synchronisation

PURPOSE: DEPENDENCIES/TEST DATA REF: PRE\_CONDITIONS:

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Reset the time on the finger oximeter  This can be achieved by removing the batteries for 30 seconds or amending the time within the settings |

|  |
| --- |
| Expected |
| Time reset to incorrect on the finger oximeter |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Take a reading via the finger Oximeter and a screen is displayed to provide instructions for reseting the time via the finger oximeter |

|  |
| --- |
| Expected |
| Time and date against the finger oximeter reset |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Retake the reading via the finger oximeter and upload the reading to MediPi  Check the reading and the date and time is correct |

|  |
| --- |
| Expected |
| Reading taken displayed correctly in MediPi against the current date and time |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| Transmit the reading to the clinical web application |

|  |
| --- |
| Expected |
| Reading transmitted to the clinical web application and displayed against the correct date and time |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# MediPi Time Synchronisation

PURPOSE: DEPENDENCIES/TEST DATA REF: PRE\_CONDITIONS:

**Test Data/Pre-Conditions**

**Testing Notes**

**Design Steps**

**Step 1**

|  |
| --- |
| Description |
| Access the Raspberry Pi desktop by either typing startx from the command prompt or by logging onto the admin section of MediPi and selecting the top toolbar. |

|  |
| --- |
| Expected |
| Raspberry Pi desktop displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 2**

|  |
| --- |
| Description |
| Turn off the WIFI so the Raspberry PI has no connection to the internet |

|  |
| --- |
| Expected |
| WIFI switched off on the Raspberry Pi |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 3**

|  |
| --- |
| Description |
| Reboot the Raspberry Pi and logon to MediPi |

|  |
| --- |
| Expected |
| Raspberry rebooted and the MediPi logon screen displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 4**

|  |
| --- |
| Description |
| Confirm at the bottom of the screen that a message is displayed - Clock not synchronised, please wait |

|  |
| --- |
| Expected |
| Warning message displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 5**

|  |
| --- |
| Description |
| Attempt to take a reading via the following tiles -  Oximeter  Scale  Blood Pressure  Thermometer  Heart Questionnaire  Scheduler |

|  |
| --- |
| Expected |
| Unable to record a reading via any of the tiles due to the time not being synchronised |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 6**

|  |
| --- |
| Description |
| Close MediPi and reboot Raspberry Pi |

|  |
| --- |
| Expected |
| Raspberry Pi restarted |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 7**

|  |
| --- |
| Description |
| Access the Raspberry Pi desktop by either typing startx from the command prompt or by logging onto the admin section of MediPi and selecting the top toolbar. |

|  |
| --- |
| Expected |
| Raspberry Pi desktop displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 8**

|  |
| --- |
| Description |
| Turn on the WIFI so the Raspberry PI has a connection to the internet |

|  |
| --- |
| Expected |
| WIFI switched on, on Raspberry Pi |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 9**

|  |
| --- |
| Description |
| Reboot the Raspberry Pi and logon to MediPi |

|  |
| --- |
| Expected |
| Raspberry rebooted and the MediPi logon screen displayed |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 10**

|  |
| --- |
| Description |
| Confirm that there is no longer a message at the bottom of the screen - Clock not synchronised, please wait  Date and time matches current time |

|  |
| --- |
| Expected |
| Message not displayed  Date and time correct as per current time |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 11**

|  |
| --- |
| Description |
| Record a couple of readings and transmit the data via MediPi |

|  |
| --- |
| Expected |
| Readings taken successfully and transmitted |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

**Step 12**

|  |
| --- |
| Description |
| Logon on to the Web Application and confirm the correct readings are populated against the correct date and time |

|  |
| --- |
| Expected |
| Correct readings recorded against current date and time |

|  |
| --- |
| Actual Result - Please include details of messages generated, screenshots and NHS Numbers used |
|  |

# Runs

## Run ID : 28741 - Fast\_Run\_10-19\_10-58-32

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28741 |
| Configuration ID: | 14371 | Run Name: | Fast\_Run\_10-19\_10-58-32 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 19/10/17 | Target Cycle: |  |
| Exec Time: | 10:58:32 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_1.0.15-PILOT-20171003-1 (Direct Message Functional Change) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28741]Heart Questionnaire - Transmit | Test: Test Name: | Heart Questionnaire - Transmit |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28740 - Fast\_Run\_10-19\_10-58-28

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28740 |
| Configuration ID: | 14405 | Run Name: | Fast\_Run\_10-19\_10-58-28 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 19/10/17 | Target Cycle: |  |
| Exec Time: | 10:58:28 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_1.0.15-PILOT-20171003-1 (Direct Message Functional Change) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28740]Diabetes Questionnaire - Transmit | Test: Test Name: | Diabetes Questionnaire - Transmit |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28739 - Fast\_Run\_10-19\_10-58-25

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28739 |
| Configuration ID: | 14404 | Run Name: | Fast\_Run\_10-19\_10-58-25 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 19/10/17 | Target Cycle: |  |
| Exec Time: | 10:58:25 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_1.0.15-PILOT-20171003-1 (Direct Message Functional Change) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28739]COPD Questionnaire - Transmit | Test: Test Name: | COPD Questionnaire - Transmit |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28738 - Fast\_Run\_10-19\_10-58-23

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28738 |
| Configuration ID: | 14370 | Run Name: | Fast\_Run\_10-19\_10-58-23 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 19/10/17 | Target Cycle: |  |
| Exec Time: | 10:58:23 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_1.0.15-PILOT-20171003-1 (Direct Message Functional Change) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28738]Blood Pressure - Transmit | Test: Test Name: | Blood Pressure - Transmit |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28737 - Fast\_Run\_10-19\_10-58-20

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28737 |
| Configuration ID: | 14359 | Run Name: | Fast\_Run\_10-19\_10-58-20 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 19/10/17 | Target Cycle: |  |
| Exec Time: | 10:58:20 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_1.0.15-PILOT-20171003-1 (Direct Message Functional Change) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28737]Blood Pressure - Record | Test: Test Name: | Blood Pressure - Record |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28736 - Fast\_Run\_10-19\_10-58-16

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28736 |
| Configuration ID: | 14382 | Run Name: | Fast\_Run\_10-19\_10-58-16 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 19/10/17 | Target Cycle: |  |
| Exec Time: | 10:58:16 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_1.0.15-PILOT-20171003-1 (Direct Message Functional Change) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28736]Thermometer - Transmit | Test: Test Name: | Thermometer - Transmit |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28735 - Fast\_Run\_10-19\_10-58-14

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28735 |
| Configuration ID: | 14381 | Run Name: | Fast\_Run\_10-19\_10-58-14 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 19/10/17 | Target Cycle: |  |
| Exec Time: | 10:58:14 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_1.0.15-PILOT-20171003-1 (Direct Message Functional Change) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28735]Thermometer - Record | Test: Test Name: | Thermometer - Record |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28734 - Fast\_Run\_10-19\_10-58-4

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28734 |
| Configuration ID: | 14368 | Run Name: | Fast\_Run\_10-19\_10-58-4 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 19/10/17 | Target Cycle: |  |
| Exec Time: | 10:58:04 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_1.0.15-PILOT-20171003-1 (Direct Message Functional Change) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28734]Finger Oximeter - Transmit | Test: Test Name: | Finger Oximeter - Transmit |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28733 - Fast\_Run\_10-19\_10-58-0

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28733 |
| Configuration ID: | 14353 | Run Name: | Fast\_Run\_10-19\_10-58-0 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 19/10/17 | Target Cycle: |  |
| Exec Time: | 10:58:00 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_1.0.15-PILOT-20171003-1 (Direct Message Functional Change) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28733]Finger Oximeter - Record | Test: Test Name: | Finger Oximeter - Record |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28728 - Fast\_Run\_9-21\_9-40-47

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28728 |
| Configuration ID: | 14353 | Run Name: | Fast\_Run\_9-21\_9-40-47 |
| Draft Run: | N | State: |  |
| Duration: | 8 | Status: | Not Completed |
| Exec Date: | 19/10/17 | Target Cycle: |  |
| Exec Time: | 10:40:44 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_1.0.15-PILOT-20170911-1 (Bluetooth Interface Engine) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28728]Finger Oximeter - Record | Test: Test Name: | Finger Oximeter - Record |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28727 - Fast\_Run\_9-20\_15-12-26

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28727 |
| Configuration ID: | 14353 | Run Name: | Fast\_Run\_9-20\_15-12-26 |
| Draft Run: | N | State: |  |
| Duration: | 336 | Status: | Not Completed |
| Exec Date: | 20/09/17 | Target Cycle: |  |
| Exec Time: | 15:18:07 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_1.0.15-PILOT-20170911-1 (Bluetooth Interface Engine) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28727]Finger Oximeter - Record | Test: Test Name: | Finger Oximeter - Record |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28726 - Fast\_Run\_9-20\_15-12-18

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28726 |
| Configuration ID: | 14368 | Run Name: | Fast\_Run\_9-20\_15-12-18 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 20/09/17 | Target Cycle: |  |
| Exec Time: | 15:12:18 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_1.0.15-PILOT-20170911-1 (Bluetooth Interface Engine) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28726]Finger Oximeter - Transmit | Test: Test Name: | Finger Oximeter - Transmit |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28725 - Fast\_Run\_9-20\_15-12-17

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28725 |
| Configuration ID: | 14346 | Run Name: | Fast\_Run\_9-20\_15-12-17 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 20/09/17 | Target Cycle: |  |
| Exec Time: | 15:12:17 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_1.0.15-PILOT-20170911-1 (Bluetooth Interface Engine) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28725]Finger Oximeter - General Screen | Test: Test Name: | Finger Oximeter - General Screen |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28724 - Fast\_Run\_9-20\_15-12-15

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28724 |
| Configuration ID: | 14385 | Run Name: | Fast\_Run\_9-20\_15-12-15 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 20/09/17 | Target Cycle: |  |
| Exec Time: | 15:12:15 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_1.0.15-PILOT-20170911-1 (Bluetooth Interface Engine) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28724]Temperature Threshold | Test: Test Name: | Temperature Threshold |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28723 - Fast\_Run\_9-20\_15-12-1

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28723 |
| Configuration ID: | 14386 | Run Name: | Fast\_Run\_9-20\_15-12-1 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 20/09/17 | Target Cycle: |  |
| Exec Time: | 15:12:01 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_1.0.15-PILOT-20170911-1 (Bluetooth Interface Engine) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28723]Pulse Rate Threshold | Test: Test Name: | Pulse Rate Threshold |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28722 - Fast\_Run\_9-20\_15-12-0

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28722 |
| Configuration ID: | 14387 | Run Name: | Fast\_Run\_9-20\_15-12-0 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 20/09/17 | Target Cycle: |  |
| Exec Time: | 15:12:00 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_1.0.15-PILOT-20170911-1 (Bluetooth Interface Engine) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28722]Oxygen Saturation Threshold | Test: Test Name: | Oxygen Saturation Threshold |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28721 - Fast\_Run\_9-20\_15-11-59

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28721 |
| Configuration ID: | 14396 | Run Name: | Fast\_Run\_9-20\_15-11-59 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 20/09/17 | Target Cycle: |  |
| Exec Time: | 15:11:59 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_1.0.15-PILOT-20170911-1 (Bluetooth Interface Engine) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28721]Blood Pressure Threshold | Test: Test Name: | Blood Pressure Threshold |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28720 - Fast\_Run\_9-20\_15-11-58

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28720 |
| Configuration ID: | 14360 | Run Name: | Fast\_Run\_9-20\_15-11-58 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 20/09/17 | Target Cycle: |  |
| Exec Time: | 15:11:58 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_1.0.15-PILOT-20170911-1 (Bluetooth Interface Engine) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28720]Scheduler - Transmit Summary Screen | Test: Test Name: | Scheduler - Transmit Summary Screen |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28719 - Fast\_Run\_9-20\_15-11-57

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28719 |
| Configuration ID: | 14374 | Run Name: | Fast\_Run\_9-20\_15-11-57 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 20/09/17 | Target Cycle: |  |
| Exec Time: | 15:11:57 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_1.0.15-PILOT-20170911-1 (Bluetooth Interface Engine) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28719]Scheduler - Transmit (Individual) | Test: Test Name: | Scheduler - Transmit (Individual) |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28718 - Fast\_Run\_9-20\_15-11-56

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28718 |
| Configuration ID: | 14371 | Run Name: | Fast\_Run\_9-20\_15-11-56 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 20/09/17 | Target Cycle: |  |
| Exec Time: | 15:11:56 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_1.0.15-PILOT-20170911-1 (Bluetooth Interface Engine) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28718]Heart Questionnaire - Transmit | Test: Test Name: | Heart Questionnaire - Transmit |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28717 - Fast\_Run\_9-20\_15-11-54

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28717 |
| Configuration ID: | 14349 | Run Name: | Fast\_Run\_9-20\_15-11-54 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 20/09/17 | Target Cycle: |  |
| Exec Time: | 15:11:54 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_1.0.15-PILOT-20170911-1 (Bluetooth Interface Engine) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28717]Heart Failure Questionnaire | Test: Test Name: | Heart Failure Questionnaire |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28716 - Fast\_Run\_9-20\_15-11-35

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28716 |
| Configuration ID: | 14405 | Run Name: | Fast\_Run\_9-20\_15-11-35 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 20/09/17 | Target Cycle: |  |
| Exec Time: | 15:11:35 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_1.0.15-PILOT-20170911-1 (Bluetooth Interface Engine) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28716]Diabetes Questionnaire - Transmit | Test: Test Name: | Diabetes Questionnaire - Transmit |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28715 - Fast\_Run\_9-20\_15-11-21

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28715 |
| Configuration ID: | 14402 | Run Name: | Fast\_Run\_9-20\_15-11-21 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 20/09/17 | Target Cycle: |  |
| Exec Time: | 15:11:21 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_1.0.15-PILOT-20170911-1 (Bluetooth Interface Engine) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28715]Diabetes Questionnaire | Test: Test Name: | Diabetes Questionnaire |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28714 - Fast\_Run\_9-20\_15-11-19

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28714 |
| Configuration ID: | 14397 | Run Name: | Fast\_Run\_9-20\_15-11-19 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 20/09/17 | Target Cycle: |  |
| Exec Time: | 15:11:19 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_1.0.15-PILOT-20170911-1 (Bluetooth Interface Engine) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28714]COPD Questionnaire Alerts | Test: Test Name: | COPD Questionnaire Alerts |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28713 - Fast\_Run\_9-20\_15-11-17

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28713 |
| Configuration ID: | 14404 | Run Name: | Fast\_Run\_9-20\_15-11-17 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 20/09/17 | Target Cycle: |  |
| Exec Time: | 15:11:17 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_1.0.15-PILOT-20170911-1 (Bluetooth Interface Engine) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28713]COPD Questionnaire - Transmit | Test: Test Name: | COPD Questionnaire - Transmit |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28712 - Fast\_Run\_9-20\_15-11-16

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28712 |
| Configuration ID: | 14403 | Run Name: | Fast\_Run\_9-20\_15-11-16 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 20/09/17 | Target Cycle: |  |
| Exec Time: | 15:11:16 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_1.0.15-PILOT-20170911-1 (Bluetooth Interface Engine) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28712]COPD Questionnaire | Test: Test Name: | COPD Questionnaire |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28711 - Fast\_Run\_9-20\_15-11-14

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28711 |
| Configuration ID: | 14370 | Run Name: | Fast\_Run\_9-20\_15-11-14 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 20/09/17 | Target Cycle: |  |
| Exec Time: | 15:11:14 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_1.0.15-PILOT-20170911-1 (Bluetooth Interface Engine) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28711]Blood Pressure - Transmit | Test: Test Name: | Blood Pressure - Transmit |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28710 - Fast\_Run\_9-20\_15-11-13

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28710 |
| Configuration ID: | 14359 | Run Name: | Fast\_Run\_9-20\_15-11-13 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 20/09/17 | Target Cycle: |  |
| Exec Time: | 15:11:13 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_1.0.15-PILOT-20170911-1 (Bluetooth Interface Engine) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28710]Blood Pressure - Record | Test: Test Name: | Blood Pressure - Record |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28709 - Fast\_Run\_9-20\_15-10-52

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28709 |
| Configuration ID: | 14348 | Run Name: | Fast\_Run\_9-20\_15-10-52 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 20/09/17 | Target Cycle: |  |
| Exec Time: | 15:10:52 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_1.0.15-PILOT-20170911-1 (Bluetooth Interface Engine) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28709]Blood Pressure - General Screen | Test: Test Name: | Blood Pressure - General Screen |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28708 - Fast\_Run\_9-20\_15-10-48

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28708 |
| Configuration ID: | 14382 | Run Name: | Fast\_Run\_9-20\_15-10-48 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 20/09/17 | Target Cycle: |  |
| Exec Time: | 15:10:48 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_1.0.15-PILOT-20170911-1 (Bluetooth Interface Engine) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28708]Thermometer - Transmit | Test: Test Name: | Thermometer - Transmit |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28707 - Fast\_Run\_9-20\_15-10-46

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28707 |
| Configuration ID: | 14381 | Run Name: | Fast\_Run\_9-20\_15-10-46 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 20/09/17 | Target Cycle: |  |
| Exec Time: | 15:10:46 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_1.0.15-PILOT-20170911-1 (Bluetooth Interface Engine) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28707]Thermometer - Record | Test: Test Name: | Thermometer - Record |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28706 - Fast\_Run\_9-20\_15-10-43

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28706 |
| Configuration ID: | 14380 | Run Name: | Fast\_Run\_9-20\_15-10-43 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 20/09/17 | Target Cycle: |  |
| Exec Time: | 15:10:43 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_1.0.15-PILOT-20170911-1 (Bluetooth Interface Engine) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28706]Thermometer - General Screen | Test: Test Name: | Thermometer - General Screen |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28535 - Run\_7-25\_12-48-27

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 28535 |
| Configuration ID: | 14389 | Run Name: | Run\_7-25\_12-48-27 |
| Draft Run: | N | State: |  |
| Duration: | 48 | Status: | Passed |
| Exec Date: | 25/07/17 | Target Cycle: |  |
| Exec Time: | 12:49:19 | Test Instance: |  |
| Host: | RIRO-PC | Test Set: Name: | Clinical Web Application (N3) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28535]Caseload View - Question Mark | Test: Test Name: | Caseload View - Question Mark |
| OS Build Number: | Build 7601 | Tester: | riro(Richard Robinson) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28529 - Run\_7-24\_17-56-45

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 28529 |
| Configuration ID: | 14393 | Run Name: | Run\_7-24\_17-56-45 |
| Draft Run: | N | State: |  |
| Duration: | 1007 | Status: | Passed |
| Exec Date: | 24/07/17 | Target Cycle: |  |
| Exec Time: | 18:14:31 | Test Instance: |  |
| Host: | RIRO-PC | Test Set: Name: | Clinical Web Application (N3) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28529]Caseload View - Sad Face | Test: Test Name: | Caseload View - Sad Face |
| OS Build Number: | Build 7601 | Tester: | riro(Richard Robinson) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28528 - Run\_7-24\_17-47-7

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 28528 |
| Configuration ID: | 14378 | Run Name: | Run\_7-24\_17-47-7 |
| Draft Run: | N | State: |  |
| Duration: | 287 | Status: | Passed |
| Exec Date: | 24/07/17 | Target Cycle: |  |
| Exec Time: | 17:52:39 | Test Instance: |  |
| Host: | RIRO-PC | Test Set: Name: | Clinical Web Application (N3) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28528]Caseload View - Happy Face | Test: Test Name: | Caseload View - Happy Face |
| OS Build Number: | Build 7601 | Tester: | riro(Richard Robinson) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28527 - Run\_7-24\_17-30-33

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 28527 |
| Configuration ID: | 14385 | Run Name: | Run\_7-24\_17-30-33 |
| Draft Run: | N | State: |  |
| Duration: | 979 | Status: | Passed |
| Exec Date: | 24/07/17 | Target Cycle: |  |
| Exec Time: | 17:47:31 | Test Instance: |  |
| Host: | RIRO-PC | Test Set: Name: | Clinical Web Application (N3) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28527]Temperature Threshold | Test: Test Name: | Temperature Threshold |
| OS Build Number: | Build 7601 | Tester: | riro(Richard Robinson) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28526 - Run\_7-24\_17-19-12

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 28526 |
| Configuration ID: | 14373 | Run Name: | Run\_7-24\_17-19-12 |
| Draft Run: | N | State: |  |
| Duration: | 667 | Status: | Passed |
| Exec Date: | 24/07/17 | Target Cycle: |  |
| Exec Time: | 17:30:43 | Test Instance: |  |
| Host: | RIRO-PC | Test Set: Name: | Clinical Web Application (N3) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28526]Scheduler - Transmit (All Data) | Test: Test Name: | Scheduler - Transmit (All Data) |
| OS Build Number: | Build 7601 | Tester: | riro(Richard Robinson) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28525 - Run\_7-24\_17-13-54

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 28525 |
| Configuration ID: | 14469 | Run Name: | Run\_7-24\_17-13-54 |
| Draft Run: | N | State: |  |
| Duration: | 243 | Status: | Passed |
| Exec Date: | 24/07/17 | Target Cycle: |  |
| Exec Time: | 17:18:10 | Test Instance: |  |
| Host: | RIRO-PC | Test Set: Name: | Clinical Web Application (N3) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28525]Logon & Logoff | Test: Test Name: | Logon & Logoff |
| OS Build Number: | Build 7601 | Tester: | riro(Richard Robinson) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28523 - Run\_7-24\_17-13-38

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 28523 |
| Configuration ID: | 14469 | Run Name: | Run\_7-24\_17-13-38 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: |  |
| Exec Date: | 24/07/17 | Target Cycle: |  |
| Exec Time: | 17:13:45 | Test Instance: |  |
| Host: |  | Test Set: Name: | Clinical Web Application (N3) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28523]Logon & Logoff | Test: Test Name: | Logon & Logoff |
| OS Build Number: | Build 7601 | Tester: | riro(Richard Robinson) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28522 - Run\_7-24\_17-13-35

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 28522 |
| Configuration ID: | 14389 | Run Name: | Run\_7-24\_17-13-35 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: |  |
| Exec Date: | 24/07/17 | Target Cycle: |  |
| Exec Time: | 17:13:42 | Test Instance: |  |
| Host: |  | Test Set: Name: | Clinical Web Application (N3) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28522]Caseload View - Question Mark | Test: Test Name: | Caseload View - Question Mark |
| OS Build Number: | Build 7601 | Tester: | riro(Richard Robinson) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28521 - Run\_7-24\_17-10-22

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 28521 |
| Configuration ID: | 14393 | Run Name: | Run\_7-24\_17-10-22 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: |  |
| Exec Date: | 24/07/17 | Target Cycle: |  |
| Exec Time: | 17:10:29 | Test Instance: |  |
| Host: |  | Test Set: Name: | Clinical Web Application (N3) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28521]Caseload View - Sad Face | Test: Test Name: | Caseload View - Sad Face |
| OS Build Number: | Build 7601 | Tester: | riro(Richard Robinson) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28519 - Run\_7-24\_17-8-25

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 28519 |
| Configuration ID: | 14385 | Run Name: | Run\_7-24\_17-8-25 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: |  |
| Exec Date: | 24/07/17 | Target Cycle: |  |
| Exec Time: | 17:08:30 | Test Instance: |  |
| Host: |  | Test Set: Name: | Clinical Web Application (N3) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28519]Temperature Threshold | Test: Test Name: | Temperature Threshold |
| OS Build Number: | Build 7601 | Tester: | riro(Richard Robinson) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28163 - Run\_6-21\_14-19-35

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 28163 |
| Configuration ID: | 14811 | Run Name: | Run\_6-21\_14-19-35 |
| Draft Run: | N | State: |  |
| Duration: | 346 | Status: | Passed |
| Exec Date: | 21/06/17 | Target Cycle: |  |
| Exec Time: | 14:25:21 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15-PILOT-20170530-1 (Functional changes as requested by Hertfordshire Trust) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28163]MediPi Clinical Application - Weight thresholds changed to be simple high & low | Test: Test Name: | MediPi Clinical Application - Weight thresholds changed to be simple high & low |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28162 - Run\_6-21\_14-19-24

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 28162 |
| Configuration ID: | 14808 | Run Name: | Run\_6-21\_14-19-24 |
| Draft Run: | N | State: |  |
| Duration: | 7 | Status: | Passed |
| Exec Date: | 21/06/17 | Target Cycle: |  |
| Exec Time: | 14:19:31 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15-PILOT-20170530-1 (Functional changes as requested by Hertfordshire Trust) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28162]MediPi Application - Timeout increased to 15 minutes | Test: Test Name: | MediPi Application - Timeout increased to 15 minutes |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28161 - Run\_6-21\_14-18-6

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 28161 |
| Configuration ID: | 14809 | Run Name: | Run\_6-21\_14-18-6 |
| Draft Run: | N | State: |  |
| Duration: | 68 | Status: | Passed |
| Exec Date: | 21/06/17 | Target Cycle: |  |
| Exec Time: | 14:19:14 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15-PILOT-20170530-1 (Functional changes as requested by Hertfordshire Trust) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28161]MediPi Application - Removal of all tiles on the home screen except for the readings tile | Test: Test Name: | MediPi Application - Removal of all tiles on the home screen except for the readings tile |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28160 - Run\_6-21\_14-17-54

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 28160 |
| Configuration ID: | 14810 | Run Name: | Run\_6-21\_14-17-54 |
| Draft Run: | N | State: |  |
| Duration: | 7 | Status: | Passed |
| Exec Date: | 21/06/17 | Target Cycle: |  |
| Exec Time: | 14:18:02 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15-PILOT-20170530-1 (Functional changes as requested by Hertfordshire Trust) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28160]MediPi Application - Order of readings changed via schedule | Test: Test Name: | MediPi Application - Order of readings changed via schedule |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28159 - Run\_6-21\_14-17-27

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 28159 |
| Configuration ID: | 14807 | Run Name: | Run\_6-21\_14-17-27 |
| Draft Run: | N | State: |  |
| Duration: | 15 | Status: | Passed |
| Exec Date: | 21/06/17 | Target Cycle: |  |
| Exec Time: | 14:17:42 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15-PILOT-20170530-1 (Functional changes as requested by Hertfordshire Trust) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28159]MediPi Application - Removal of PID on main MediPi Screen before authentication | Test: Test Name: | MediPi Application - Removal of PID on main MediPi Screen before authentication |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28055 - Fast\_Run\_5-9\_12-55-39

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28055 |
| Configuration ID: | 14365 | Run Name: | Fast\_Run\_5-9\_12-55-39 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 09/05/17 | Target Cycle: |  |
| Exec Time: | 12:55:39 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_v1.0.8-PILOT-20160907-1 (MediPi App/USB Connected Apps) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28055]MediPi Patient Banner - Patient Name | Test: Test Name: | MediPi Patient Banner - Patient Name |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28054 - Fast\_Run\_5-9\_12-55-36

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28054 |
| Configuration ID: | 14364 | Run Name: | Fast\_Run\_5-9\_12-55-36 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 09/05/17 | Target Cycle: |  |
| Exec Time: | 12:55:36 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_v1.0.8-PILOT-20160907-1 (MediPi App/USB Connected Apps) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28054]Fatal Error Screen | Test: Test Name: | Fatal Error Screen |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28053 - Fast\_Run\_5-9\_12-55-36

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28053 |
| Configuration ID: | 14352 | Run Name: | Fast\_Run\_5-9\_12-55-36 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 09/05/17 | Target Cycle: |  |
| Exec Time: | 12:55:36 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_v1.0.8-PILOT-20160907-1 (MediPi App/USB Connected Apps) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28053]Transmitter - General Screen | Test: Test Name: | Transmitter - General Screen |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28052 - Fast\_Run\_5-9\_12-55-35

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28052 |
| Configuration ID: | 14366 | Run Name: | Fast\_Run\_5-9\_12-55-35 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 09/05/17 | Target Cycle: |  |
| Exec Time: | 12:55:35 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_v1.0.8-PILOT-20160907-1 (MediPi App/USB Connected Apps) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28052]Clinician Messages - Polling Warning | Test: Test Name: | Clinician Messages - Polling Warning |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28051 - Fast\_Run\_5-9\_12-55-34

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28051 |
| Configuration ID: | 14360 | Run Name: | Fast\_Run\_5-9\_12-55-34 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 09/05/17 | Target Cycle: |  |
| Exec Time: | 12:55:34 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_v1.0.8-PILOT-20160907-1 (MediPi App/USB Connected Apps) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28051]Scheduler - Transmit Summary Screen | Test: Test Name: | Scheduler - Transmit Summary Screen |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28050 - Fast\_Run\_5-9\_12-55-33

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28050 |
| Configuration ID: | 14358 | Run Name: | Fast\_Run\_5-9\_12-55-33 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 09/05/17 | Target Cycle: |  |
| Exec Time: | 12:55:33 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_v1.0.8-PILOT-20160907-1 (MediPi App/USB Connected Apps) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28050]Diagnostic Scale - Record | Test: Test Name: | Diagnostic Scale - Record |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28049 - Fast\_Run\_5-9\_12-55-33

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28049 |
| Configuration ID: | 14353 | Run Name: | Fast\_Run\_5-9\_12-55-33 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 09/05/17 | Target Cycle: |  |
| Exec Time: | 12:55:33 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_v1.0.8-PILOT-20160907-1 (MediPi App/USB Connected Apps) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28049]Finger Oximeter - Record | Test: Test Name: | Finger Oximeter - Record |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28048 - Fast\_Run\_5-9\_12-55-32

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28048 |
| Configuration ID: | 14346 | Run Name: | Fast\_Run\_5-9\_12-55-32 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 09/05/17 | Target Cycle: |  |
| Exec Time: | 12:55:32 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_v1.0.8-PILOT-20160907-1 (MediPi App/USB Connected Apps) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28048]Finger Oximeter - General Screen | Test: Test Name: | Finger Oximeter - General Screen |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28047 - Fast\_Run\_5-9\_12-55-31

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28047 |
| Configuration ID: | 14362 | Run Name: | Fast\_Run\_5-9\_12-55-31 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 09/05/17 | Target Cycle: |  |
| Exec Time: | 12:55:31 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_v1.0.8-PILOT-20160907-1 (MediPi App/USB Connected Apps) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28047]Tile Configuration | Test: Test Name: | Tile Configuration |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28046 - Fast\_Run\_5-9\_12-55-26

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28046 |
| Configuration ID: | 14418 | Run Name: | Fast\_Run\_5-9\_12-55-26 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 09/05/17 | Target Cycle: |  |
| Exec Time: | 12:55:26 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_v1.0.14 PILOT-20170127-1 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28046]Finger Oximeter - Time Synchronisation | Test: Test Name: | Finger Oximeter - Time Synchronisation |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28045 - Fast\_Run\_5-9\_12-55-23

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28045 |
| Configuration ID: | 14352 | Run Name: | Fast\_Run\_5-9\_12-55-23 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 09/05/17 | Target Cycle: |  |
| Exec Time: | 12:55:23 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_1.0.9 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28045]Transmitter - General Screen | Test: Test Name: | Transmitter - General Screen |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28044 - Fast\_Run\_5-9\_12-55-21

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28044 |
| Configuration ID: | 14351 | Run Name: | Fast\_Run\_5-9\_12-55-21 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 09/05/17 | Target Cycle: |  |
| Exec Time: | 12:55:21 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_1.0.9 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28044]Clinician Messages - General Screen | Test: Test Name: | Clinician Messages - General Screen |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28043 - Fast\_Run\_5-9\_12-55-20

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28043 |
| Configuration ID: | 14358 | Run Name: | Fast\_Run\_5-9\_12-55-20 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 09/05/17 | Target Cycle: |  |
| Exec Time: | 12:55:20 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_1.0.9 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28043]Diagnostic Scale - Record | Test: Test Name: | Diagnostic Scale - Record |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28042 - Fast\_Run\_5-9\_12-55-18

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28042 |
| Configuration ID: | 14380 | Run Name: | Fast\_Run\_5-9\_12-55-18 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 09/05/17 | Target Cycle: |  |
| Exec Time: | 12:55:19 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_1.0.9 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28042]Thermometer - General Screen | Test: Test Name: | Thermometer - General Screen |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28041 - Fast\_Run\_5-9\_12-55-17

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28041 |
| Configuration ID: | 14346 | Run Name: | Fast\_Run\_5-9\_12-55-17 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 09/05/17 | Target Cycle: |  |
| Exec Time: | 12:55:17 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_1.0.9 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28041]Finger Oximeter - General Screen | Test: Test Name: | Finger Oximeter - General Screen |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28040 - Fast\_Run\_5-9\_12-55-16

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28040 |
| Configuration ID: | 14362 | Run Name: | Fast\_Run\_5-9\_12-55-16 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 09/05/17 | Target Cycle: |  |
| Exec Time: | 12:55:16 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_1.0.9 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28040]Tile Configuration | Test: Test Name: | Tile Configuration |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28039 - Fast\_Run\_5-9\_12-55-12

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28039 |
| Configuration ID: | 14361 | Run Name: | Fast\_Run\_5-9\_12-55-12 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 09/05/17 | Target Cycle: |  |
| Exec Time: | 12:55:12 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_1.0.9 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28039]Homepage - General Screen | Test: Test Name: | Homepage - General Screen |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28038 - Fast\_Run\_5-9\_12-55-8

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28038 |
| Configuration ID: | 14395 | Run Name: | Fast\_Run\_5-9\_12-55-8 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 09/05/17 | Target Cycle: |  |
| Exec Time: | 12:55:08 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi v1.0.14 PILOT - 20170125 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28038]Weight Threshold | Test: Test Name: | Weight Threshold |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28037 - Fast\_Run\_5-9\_12-55-6

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28037 |
| Configuration ID: | 14396 | Run Name: | Fast\_Run\_5-9\_12-55-6 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 09/05/17 | Target Cycle: |  |
| Exec Time: | 12:55:06 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi v1.0.14 PILOT - 20170125 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28037]Blood Pressure Threshold | Test: Test Name: | Blood Pressure Threshold |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28036 - Fast\_Run\_5-9\_12-55-3

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28036 |
| Configuration ID: | 14414 | Run Name: | Fast\_Run\_5-9\_12-55-3 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 09/05/17 | Target Cycle: |  |
| Exec Time: | 12:55:03 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi v1.0.14 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28036]Admin Configuration Screen | Test: Test Name: | Admin Configuration Screen |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28035 - Fast\_Run\_5-9\_12-54-59

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28035 |
| Configuration ID: | 14353 | Run Name: | Fast\_Run\_5-9\_12-54-59 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 09/05/17 | Target Cycle: |  |
| Exec Time: | 12:54:59 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi V1.0.11 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28035]Finger Oximeter - Record | Test: Test Name: | Finger Oximeter - Record |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28034 - Fast\_Run\_5-9\_12-54-55

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28034 |
| Configuration ID: | 14395 | Run Name: | Fast\_Run\_5-9\_12-54-55 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 09/05/17 | Target Cycle: |  |
| Exec Time: | 12:54:55 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28034]Weight Threshold | Test: Test Name: | Weight Threshold |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28033 - Fast\_Run\_5-9\_12-54-51

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28033 |
| Configuration ID: | 14384 | Run Name: | Fast\_Run\_5-9\_12-54-51 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 09/05/17 | Target Cycle: |  |
| Exec Time: | 12:54:51 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28033]Threshold Input Value Validation | Test: Test Name: | Threshold Input Value Validation |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28032 - Fast\_Run\_5-9\_12-54-50

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28032 |
| Configuration ID: | 14385 | Run Name: | Fast\_Run\_5-9\_12-54-50 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 09/05/17 | Target Cycle: |  |
| Exec Time: | 12:54:50 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28032]Temperature Threshold | Test: Test Name: | Temperature Threshold |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28031 - Fast\_Run\_5-9\_12-54-48

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28031 |
| Configuration ID: | 14386 | Run Name: | Fast\_Run\_5-9\_12-54-48 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 09/05/17 | Target Cycle: |  |
| Exec Time: | 12:54:48 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28031]Pulse Rate Threshold | Test: Test Name: | Pulse Rate Threshold |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28030 - Fast\_Run\_5-9\_12-54-47

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28030 |
| Configuration ID: | 14387 | Run Name: | Fast\_Run\_5-9\_12-54-47 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 09/05/17 | Target Cycle: |  |
| Exec Time: | 12:54:47 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28030]Oxygen Saturation Threshold | Test: Test Name: | Oxygen Saturation Threshold |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28029 - Fast\_Run\_5-9\_12-54-46

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28029 |
| Configuration ID: | 14365 | Run Name: | Fast\_Run\_5-9\_12-54-46 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 09/05/17 | Target Cycle: |  |
| Exec Time: | 12:54:47 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28029]MediPi Patient Banner - Patient Name | Test: Test Name: | MediPi Patient Banner - Patient Name |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28028 - Fast\_Run\_5-9\_12-54-46

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28028 |
| Configuration ID: | 14375 | Run Name: | Fast\_Run\_5-9\_12-54-46 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 09/05/17 | Target Cycle: |  |
| Exec Time: | 12:54:46 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28028]Browser Testing - Internet Explorer | Test: Test Name: | Browser Testing - Internet Explorer |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28027 - Fast\_Run\_5-9\_12-54-45

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28027 |
| Configuration ID: | 14388 | Run Name: | Fast\_Run\_5-9\_12-54-45 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 09/05/17 | Target Cycle: |  |
| Exec Time: | 12:54:45 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28027]Heart Questionnaire - Alerts | Test: Test Name: | Heart Questionnaire - Alerts |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28026 - Fast\_Run\_5-9\_12-54-43

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28026 |
| Configuration ID: | 14398 | Run Name: | Fast\_Run\_5-9\_12-54-43 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 09/05/17 | Target Cycle: |  |
| Exec Time: | 12:54:43 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28026]Diabetes Questionnaire Alerts | Test: Test Name: | Diabetes Questionnaire Alerts |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28025 - Fast\_Run\_5-9\_12-54-42

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28025 |
| Configuration ID: | 14397 | Run Name: | Fast\_Run\_5-9\_12-54-42 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 09/05/17 | Target Cycle: |  |
| Exec Time: | 12:54:42 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28025]COPD Questionnaire Alerts | Test: Test Name: | COPD Questionnaire Alerts |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28024 - Fast\_Run\_5-9\_12-54-36

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28024 |
| Configuration ID: | 14348 | Run Name: | Fast\_Run\_5-9\_12-54-36 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 09/05/17 | Target Cycle: |  |
| Exec Time: | 12:54:36 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28024]Blood Pressure - General Screen | Test: Test Name: | Blood Pressure - General Screen |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28023 - Fast\_Run\_5-9\_12-54-33

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28023 |
| Configuration ID: | 14369 | Run Name: | Fast\_Run\_5-9\_12-54-33 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 09/05/17 | Target Cycle: |  |
| Exec Time: | 12:54:33 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28023]Diagnostic Scale - Transmit | Test: Test Name: | Diagnostic Scale - Transmit |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28022 - Fast\_Run\_5-9\_12-54-30

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 28022 |
| Configuration ID: | 14362 | Run Name: | Fast\_Run\_5-9\_12-54-30 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 09/05/17 | Target Cycle: |  |
| Exec Time: | 12:54:30 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28022]Tile Configuration | Test: Test Name: | Tile Configuration |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28021 - Run\_5-9\_11-13-28

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 28021 |
| Configuration ID: | 14391 | Run Name: | Run\_5-9\_11-13-28 |
| Draft Run: | N | State: |  |
| Duration: | 44 | Status: | Passed |
| Exec Date: | 09/05/17 | Target Cycle: |  |
| Exec Time: | 11:14:13 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28021]Caseload View - Exclamation Mark | Test: Test Name: | Caseload View - Exclamation Mark |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 28020 - Run\_5-9\_11-12-50

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 28020 |
| Configuration ID: | 14416 | Run Name: | Run\_5-9\_11-12-50 |
| Draft Run: | N | State: |  |
| Duration: | 21 | Status: | Passed |
| Exec Date: | 09/05/17 | Target Cycle: |  |
| Exec Time: | 11:13:11 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [28020]Data transmitted from multiple MediPi Devices | Test: Test Name: | Data transmitted from multiple MediPi Devices |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27844 - Fast\_Run\_4-3\_15-20-20

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 27844 |
| Configuration ID: | 14384 | Run Name: | Fast\_Run\_4-3\_15-20-20 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | N/A |
| Exec Date: | 03/04/17 | Target Cycle: |  |
| Exec Time: | 15:20:20 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27844]Threshold Input Value Validation | Test: Test Name: | Threshold Input Value Validation |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27843 - Run\_4-3\_14-46-28

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27843 |
| Configuration ID: | 14396 | Run Name: | Run\_4-3\_14-46-28 |
| Draft Run: | N | State: |  |
| Duration: | 721 | Status: | Passed |
| Exec Date: | 03/04/17 | Target Cycle: |  |
| Exec Time: | 14:58:29 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27843]Blood Pressure Threshold | Test: Test Name: | Blood Pressure Threshold |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27842 - Run\_4-3\_14-45-59

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27842 |
| Configuration ID: | 14387 | Run Name: | Run\_4-3\_14-45-59 |
| Draft Run: | N | State: |  |
| Duration: | 7 | Status: | Passed |
| Exec Date: | 03/04/17 | Target Cycle: |  |
| Exec Time: | 14:46:07 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27842]Oxygen Saturation Threshold | Test: Test Name: | Oxygen Saturation Threshold |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27841 - Run\_4-3\_14-29-50

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27841 |
| Configuration ID: | 14386 | Run Name: | Run\_4-3\_14-29-50 |
| Draft Run: | N | State: |  |
| Duration: | 956 | Status: | Passed |
| Exec Date: | 03/04/17 | Target Cycle: |  |
| Exec Time: | 14:45:47 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27841]Pulse Rate Threshold | Test: Test Name: | Pulse Rate Threshold |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27840 - Run\_4-3\_14-21-55

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27840 |
| Configuration ID: | 14385 | Run Name: | Run\_4-3\_14-21-55 |
| Draft Run: | N | State: |  |
| Duration: | 427 | Status: | Passed |
| Exec Date: | 03/04/17 | Target Cycle: |  |
| Exec Time: | 14:29:03 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27840]Temperature Threshold | Test: Test Name: | Temperature Threshold |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27838 - Fast\_Run\_4-3\_14-16-50

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 27838 |
| Configuration ID: | 14420 | Run Name: | Fast\_Run\_4-3\_14-16-50 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 03/04/17 | Target Cycle: |  |
| Exec Time: | 14:16:50 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27838]Blood Pressure Monitor - Time Synchronisation | Test: Test Name: | Blood Pressure Monitor - Time Synchronisation |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27713 - Run\_3-23\_10-40-17

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27713 |
| Configuration ID: | 14415 | Run Name: | Run\_3-23\_10-40-17 |
| Draft Run: | N | State: |  |
| Duration: | 404 | Status: | Passed |
| Exec Date: | 23/03/17 | Target Cycle: |  |
| Exec Time: | 10:47:02 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27713]MediPi Time Synchronisation | Test: Test Name: | MediPi Time Synchronisation |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27712 - Fast\_Run\_3-23\_10-36-41

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 27712 |
| Configuration ID: | 14388 | Run Name: | Fast\_Run\_3-23\_10-36-41 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | N/A |
| Exec Date: | 23/03/17 | Target Cycle: |  |
| Exec Time: | 10:36:41 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27712]Heart Questionnaire - Alerts | Test: Test Name: | Heart Questionnaire - Alerts |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27711 - Run\_3-23\_10-36-29

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27711 |
| Configuration ID: | 14388 | Run Name: | Run\_3-23\_10-36-29 |
| Draft Run: | N | State: |  |
| Duration: | 6 | Status: | Not Completed |
| Exec Date: | 23/03/17 | Target Cycle: |  |
| Exec Time: | 10:36:36 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27711]Heart Questionnaire - Alerts | Test: Test Name: | Heart Questionnaire - Alerts |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27710 - Fast\_Run\_3-23\_10-36-28

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 27710 |
| Configuration ID: | 14398 | Run Name: | Fast\_Run\_3-23\_10-36-28 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | N/A |
| Exec Date: | 23/03/17 | Target Cycle: |  |
| Exec Time: | 10:36:28 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27710]Diabetes Questionnaire Alerts | Test: Test Name: | Diabetes Questionnaire Alerts |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27709 - Run\_3-23\_10-36-16

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27709 |
| Configuration ID: | 14398 | Run Name: | Run\_3-23\_10-36-16 |
| Draft Run: | N | State: |  |
| Duration: | 6 | Status: | Not Completed |
| Exec Date: | 23/03/17 | Target Cycle: |  |
| Exec Time: | 10:36:23 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27709]Diabetes Questionnaire Alerts | Test: Test Name: | Diabetes Questionnaire Alerts |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27708 - Fast\_Run\_3-23\_10-36-15

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 27708 |
| Configuration ID: | 14397 | Run Name: | Fast\_Run\_3-23\_10-36-15 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | N/A |
| Exec Date: | 23/03/17 | Target Cycle: |  |
| Exec Time: | 10:36:15 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27708]COPD Questionnaire Alerts | Test: Test Name: | COPD Questionnaire Alerts |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27707 - Run\_3-23\_10-34-23

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27707 |
| Configuration ID: | 14418 | Run Name: | Run\_3-23\_10-34-23 |
| Draft Run: | N | State: |  |
| Duration: | 19 | Status: | Passed |
| Exec Date: | 23/03/17 | Target Cycle: |  |
| Exec Time: | 10:34:43 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27707]Finger Oximeter - Time Synchronisation | Test: Test Name: | Finger Oximeter - Time Synchronisation |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27706 - Run\_3-23\_10-32-23

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27706 |
| Configuration ID: | 14420 | Run Name: | Run\_3-23\_10-32-23 |
| Draft Run: | N | State: |  |
| Duration: | 112 | Status: | Failed |
| Exec Date: | 23/03/17 | Target Cycle: |  |
| Exec Time: | 10:34:15 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27706]Blood Pressure Monitor - Time Synchronisation | Test: Test Name: | Blood Pressure Monitor - Time Synchronisation |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27705 - Run\_3-22\_13-55-54

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27705 |
| Configuration ID: | 14378 | Run Name: | Run\_3-22\_13-55-54 |
| Draft Run: | N | State: |  |
| Duration: | 309 | Status: | Passed |
| Exec Date: | 09/05/17 | Target Cycle: |  |
| Exec Time: | 11:12:39 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27705]Caseload View - Happy Face | Test: Test Name: | Caseload View - Happy Face |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27704 - Run\_3-22\_13-53-18

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27704 |
| Configuration ID: | 14389 | Run Name: | Run\_3-22\_13-53-18 |
| Draft Run: | N | State: |  |
| Duration: | 12 | Status: | Passed |
| Exec Date: | 22/03/17 | Target Cycle: |  |
| Exec Time: | 13:53:31 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27704]Caseload View - Question Mark | Test: Test Name: | Caseload View - Question Mark |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27703 - Run\_3-22\_13-52-41

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27703 |
| Configuration ID: | 14394 | Run Name: | Run\_3-22\_13-52-41 |
| Draft Run: | N | State: |  |
| Duration: | 23 | Status: | Passed |
| Exec Date: | 22/03/17 | Target Cycle: |  |
| Exec Time: | 13:53:05 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27703]Caseload View - Question Mark Reset | Test: Test Name: | Caseload View - Question Mark Reset |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27702 - Run\_3-22\_12-19-56

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27702 |
| Configuration ID: | 14362 | Run Name: | Run\_3-22\_12-19-56 |
| Draft Run: | N | State: |  |
| Duration: | 5340 | Status: | Passed |
| Exec Date: | 22/03/17 | Target Cycle: |  |
| Exec Time: | 13:48:57 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27702]Tile Configuration | Test: Test Name: | Tile Configuration |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27701 - Run\_3-22\_11-45-52

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27701 |
| Configuration ID: | 14364 | Run Name: | Run\_3-22\_11-45-52 |
| Draft Run: | N | State: |  |
| Duration: | 275 | Status: | Passed |
| Exec Date: | 22/03/17 | Target Cycle: |  |
| Exec Time: | 11:50:29 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27701]Fatal Error Screen | Test: Test Name: | Fatal Error Screen |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27700 - Run\_3-22\_11-14-14

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27700 |
| Configuration ID: | 14406 | Run Name: | Run\_3-22\_11-14-14 |
| Draft Run: | N | State: |  |
| Duration: | 39 | Status: | Passed |
| Exec Date: | 22/03/17 | Target Cycle: |  |
| Exec Time: | 11:14:54 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27700]Scheduler Configuration | Test: Test Name: | Scheduler Configuration |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27699 - Run\_3-22\_10-11-39

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27699 |
| Configuration ID: | 14393 | Run Name: | Run\_3-22\_10-11-39 |
| Draft Run: | N | State: |  |
| Duration: | 92 | Status: | Passed |
| Exec Date: | 22/03/17 | Target Cycle: |  |
| Exec Time: | 10:13:12 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27699]Caseload View - Sad Face | Test: Test Name: | Caseload View - Sad Face |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27698 - Run\_3-22\_10-8-54

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27698 |
| Configuration ID: | 14373 | Run Name: | Run\_3-22\_10-8-54 |
| Draft Run: | N | State: |  |
| Duration: | 83 | Status: | Passed |
| Exec Date: | 22/03/17 | Target Cycle: |  |
| Exec Time: | 10:10:18 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27698]Scheduler - Transmit (All Data) | Test: Test Name: | Scheduler - Transmit (All Data) |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27697 - Run\_3-22\_10-3-12

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27697 |
| Configuration ID: | 14374 | Run Name: | Run\_3-22\_10-3-12 |
| Draft Run: | N | State: |  |
| Duration: | 338 | Status: | Passed |
| Exec Date: | 22/03/17 | Target Cycle: |  |
| Exec Time: | 10:08:51 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27697]Scheduler - Transmit (Individual) | Test: Test Name: | Scheduler - Transmit (Individual) |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27696 - Run\_3-22\_10-1-20

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27696 |
| Configuration ID: | 14360 | Run Name: | Run\_3-22\_10-1-20 |
| Draft Run: | N | State: |  |
| Duration: | 91 | Status: | Passed |
| Exec Date: | 22/03/17 | Target Cycle: |  |
| Exec Time: | 10:02:53 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27696]Scheduler - Transmit Summary Screen | Test: Test Name: | Scheduler - Transmit Summary Screen |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27695 - Run\_3-22\_9-40-49

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27695 |
| Configuration ID: | 14350 | Run Name: | Run\_3-22\_9-40-49 |
| Draft Run: | N | State: |  |
| Duration: | 1217 | Status: | Passed |
| Exec Date: | 22/03/17 | Target Cycle: |  |
| Exec Time: | 10:01:07 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27695]Scheduler - Schedule Screens | Test: Test Name: | Scheduler - Schedule Screens |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27694 - Run\_3-22\_9-23-52

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27694 |
| Configuration ID: | 14367 | Run Name: | Run\_3-22\_9-23-52 |
| Draft Run: | N | State: |  |
| Duration: | 128 | Status: | Passed |
| Exec Date: | 22/03/17 | Target Cycle: |  |
| Exec Time: | 09:26:49 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27694]Transmitter - Transmit (All Data) | Test: Test Name: | Transmitter - Transmit (All Data) |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27693 - Run\_3-22\_9-17-3

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27693 |
| Configuration ID: | 14369 | Run Name: | Run\_3-22\_9-17-3 |
| Draft Run: | N | State: |  |
| Duration: | 30 | Status: | Passed |
| Exec Date: | 22/03/17 | Target Cycle: |  |
| Exec Time: | 09:17:33 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27693]Diagnostic Scale - Transmit | Test: Test Name: | Diagnostic Scale - Transmit |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27692 - Run\_3-22\_9-15-3

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27692 |
| Configuration ID: | 14358 | Run Name: | Run\_3-22\_9-15-3 |
| Draft Run: | N | State: |  |
| Duration: | 112 | Status: | Passed |
| Exec Date: | 22/03/17 | Target Cycle: |  |
| Exec Time: | 09:16:56 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27692]Diagnostic Scale - Record | Test: Test Name: | Diagnostic Scale - Record |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27690 - Run\_3-21\_15-32-46

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27690 |
| Configuration ID: | 14352 | Run Name: | Run\_3-21\_15-32-46 |
| Draft Run: | N | State: |  |
| Duration: | 744 | Status: | Passed |
| Exec Date: | 22/03/17 | Target Cycle: |  |
| Exec Time: | 09:31:52 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27690]Transmitter - General Screen | Test: Test Name: | Transmitter - General Screen |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27689 - Run\_3-21\_15-29-19

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27689 |
| Configuration ID: | 14370 | Run Name: | Run\_3-21\_15-29-19 |
| Draft Run: | N | State: |  |
| Duration: | 81 | Status: | Passed |
| Exec Date: | 21/03/17 | Target Cycle: |  |
| Exec Time: | 15:30:41 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27689]Blood Pressure - Transmit | Test: Test Name: | Blood Pressure - Transmit |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27688 - Run\_3-21\_15-26-14

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27688 |
| Configuration ID: | 14359 | Run Name: | Run\_3-21\_15-26-14 |
| Draft Run: | N | State: |  |
| Duration: | 179 | Status: | Passed |
| Exec Date: | 21/03/17 | Target Cycle: |  |
| Exec Time: | 15:29:14 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27688]Blood Pressure - Record | Test: Test Name: | Blood Pressure - Record |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27687 - Run\_3-21\_15-11-23

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27687 |
| Configuration ID: | 14363 | Run Name: | Run\_3-21\_15-11-23 |
| Draft Run: | N | State: |  |
| Duration: | 141 | Status: | Passed |
| Exec Date: | 21/03/17 | Target Cycle: |  |
| Exec Time: | 15:13:45 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27687]Multiple Connected Devices | Test: Test Name: | Multiple Connected Devices |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27686 - Fast\_Run\_3-21\_15-11-22

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 27686 |
| Configuration ID: | 14366 | Run Name: | Fast\_Run\_3-21\_15-11-22 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | N/A |
| Exec Date: | 21/03/17 | Target Cycle: |  |
| Exec Time: | 15:11:22 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27686]Clinician Messages - Polling Warning | Test: Test Name: | Clinician Messages - Polling Warning |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27685 - Run\_3-21\_15-10-15

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27685 |
| Configuration ID: | 14366 | Run Name: | Run\_3-21\_15-10-15 |
| Draft Run: | N | State: |  |
| Duration: | 32 | Status: | Not Completed |
| Exec Date: | 21/03/17 | Target Cycle: |  |
| Exec Time: | 15:10:47 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27685]Clinician Messages - Polling Warning | Test: Test Name: | Clinician Messages - Polling Warning |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27684 - Fast\_Run\_3-21\_15-10-12

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 27684 |
| Configuration ID: | 14351 | Run Name: | Fast\_Run\_3-21\_15-10-12 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | N/A |
| Exec Date: | 21/03/17 | Target Cycle: |  |
| Exec Time: | 15:10:12 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27684]Clinician Messages - General Screen | Test: Test Name: | Clinician Messages - General Screen |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27683 - Run\_3-21\_15-8-12

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27683 |
| Configuration ID: | 14351 | Run Name: | Run\_3-21\_15-8-12 |
| Draft Run: | N | State: |  |
| Duration: | 25 | Status: | Not Completed |
| Exec Date: | 21/03/17 | Target Cycle: |  |
| Exec Time: | 15:08:38 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27683]Clinician Messages - General Screen | Test: Test Name: | Clinician Messages - General Screen |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27682 - Run\_3-21\_15-3-49

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27682 |
| Configuration ID: | 14365 | Run Name: | Run\_3-21\_15-3-49 |
| Draft Run: | N | State: |  |
| Duration: | 181 | Status: | Passed |
| Exec Date: | 21/03/17 | Target Cycle: |  |
| Exec Time: | 15:06:51 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27682]MediPi Patient Banner - Patient Name | Test: Test Name: | MediPi Patient Banner - Patient Name |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27681 - Run\_3-21\_15-2-48

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27681 |
| Configuration ID: | 14417 | Run Name: | Run\_3-21\_15-2-48 |
| Draft Run: | N | State: |  |
| Duration: | 41 | Status: | Passed |
| Exec Date: | 21/03/17 | Target Cycle: |  |
| Exec Time: | 15:03:30 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27681]Loss of WIFI Connection | Test: Test Name: | Loss of WIFI Connection |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27680 - Run\_3-21\_15-2-0

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27680 |
| Configuration ID: | 14371 | Run Name: | Run\_3-21\_15-2-0 |
| Draft Run: | N | State: |  |
| Duration: | 7 | Status: | Passed |
| Exec Date: | 21/03/17 | Target Cycle: |  |
| Exec Time: | 15:02:07 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27680]Heart Questionnaire - Transmit | Test: Test Name: | Heart Questionnaire - Transmit |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27679 - Run\_3-21\_15-1-46

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27679 |
| Configuration ID: | 14405 | Run Name: | Run\_3-21\_15-1-46 |
| Draft Run: | N | State: |  |
| Duration: | 7 | Status: | Passed |
| Exec Date: | 21/03/17 | Target Cycle: |  |
| Exec Time: | 15:01:54 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27679]Diabetes Questionnaire - Transmit | Test: Test Name: | Diabetes Questionnaire - Transmit |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27678 - Run\_3-21\_15-0-28

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27678 |
| Configuration ID: | 14397 | Run Name: | Run\_3-21\_15-0-28 |
| Draft Run: | N | State: |  |
| Duration: | 112 | Status: | Not Completed |
| Exec Date: | 21/03/17 | Target Cycle: |  |
| Exec Time: | 15:01:36 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27678]COPD Questionnaire Alerts | Test: Test Name: | COPD Questionnaire Alerts |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27677 - Run\_3-21\_15-0-2

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27677 |
| Configuration ID: | 14404 | Run Name: | Run\_3-21\_15-0-2 |
| Draft Run: | N | State: |  |
| Duration: | 17 | Status: | Passed |
| Exec Date: | 21/03/17 | Target Cycle: |  |
| Exec Time: | 15:00:20 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27677]COPD Questionnaire - Transmit | Test: Test Name: | COPD Questionnaire - Transmit |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27676 - Run\_3-21\_14-58-47

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27676 |
| Configuration ID: | 14382 | Run Name: | Run\_3-21\_14-58-47 |
| Draft Run: | N | State: |  |
| Duration: | 50 | Status: | Passed |
| Exec Date: | 21/03/17 | Target Cycle: |  |
| Exec Time: | 14:59:37 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27676]Thermometer - Transmit | Test: Test Name: | Thermometer - Transmit |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27675 - Run\_3-21\_14-56-15

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27675 |
| Configuration ID: | 14381 | Run Name: | Run\_3-21\_14-56-15 |
| Draft Run: | N | State: |  |
| Duration: | 148 | Status: | Passed |
| Exec Date: | 21/03/17 | Target Cycle: |  |
| Exec Time: | 14:58:43 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27675]Thermometer - Record | Test: Test Name: | Thermometer - Record |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27674 - Run\_3-21\_14-55-21

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27674 |
| Configuration ID: | 14368 | Run Name: | Run\_3-21\_14-55-21 |
| Draft Run: | N | State: |  |
| Duration: | 47 | Status: | Passed |
| Exec Date: | 21/03/17 | Target Cycle: |  |
| Exec Time: | 14:56:08 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27674]Finger Oximeter - Transmit | Test: Test Name: | Finger Oximeter - Transmit |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27673 - Run\_3-21\_14-43-59

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27673 |
| Configuration ID: | 14353 | Run Name: | Run\_3-21\_14-43-59 |
| Draft Run: | N | State: |  |
| Duration: | 677 | Status: | Passed |
| Exec Date: | 21/03/17 | Target Cycle: |  |
| Exec Time: | 14:55:16 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27673]Finger Oximeter - Record | Test: Test Name: | Finger Oximeter - Record |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27672 - Run\_3-21\_14-42-36

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27672 |
| Configuration ID: | 14349 | Run Name: | Run\_3-21\_14-42-36 |
| Draft Run: | N | State: |  |
| Duration: | 65 | Status: | Passed |
| Exec Date: | 21/03/17 | Target Cycle: |  |
| Exec Time: | 14:43:42 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27672]Heart Failure Questionnaire | Test: Test Name: | Heart Failure Questionnaire |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27671 - Run\_3-21\_14-41-4

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27671 |
| Configuration ID: | 14402 | Run Name: | Run\_3-21\_14-41-4 |
| Draft Run: | N | State: |  |
| Duration: | 87 | Status: | Passed |
| Exec Date: | 21/03/17 | Target Cycle: |  |
| Exec Time: | 14:42:31 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27671]Diabetes Questionnaire | Test: Test Name: | Diabetes Questionnaire |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27670 - Run\_3-21\_14-33-40

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27670 |
| Configuration ID: | 14403 | Run Name: | Run\_3-21\_14-33-40 |
| Draft Run: | N | State: |  |
| Duration: | 439 | Status: | Passed |
| Exec Date: | 21/03/17 | Target Cycle: |  |
| Exec Time: | 14:40:59 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27670]COPD Questionnaire | Test: Test Name: | COPD Questionnaire |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27669 - Run\_3-21\_14-31-47

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27669 |
| Configuration ID: | 14348 | Run Name: | Run\_3-21\_14-31-47 |
| Draft Run: | N | State: |  |
| Duration: | 73 | Status: | Passed |
| Exec Date: | 21/03/17 | Target Cycle: |  |
| Exec Time: | 14:33:00 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27669]Blood Pressure - General Screen | Test: Test Name: | Blood Pressure - General Screen |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27668 - Run\_3-21\_14-29-49

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27668 |
| Configuration ID: | 14347 | Run Name: | Run\_3-21\_14-29-49 |
| Draft Run: | N | State: |  |
| Duration: | 102 | Status: | Passed |
| Exec Date: | 21/03/17 | Target Cycle: |  |
| Exec Time: | 14:31:32 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27668]Diagnostic Scale - General Screen | Test: Test Name: | Diagnostic Scale - General Screen |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27667 - Run\_3-21\_14-25-28

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27667 |
| Configuration ID: | 14380 | Run Name: | Run\_3-21\_14-25-28 |
| Draft Run: | N | State: |  |
| Duration: | 252 | Status: | Passed |
| Exec Date: | 21/03/17 | Target Cycle: |  |
| Exec Time: | 14:29:40 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27667]Thermometer - General Screen | Test: Test Name: | Thermometer - General Screen |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27666 - Run\_3-21\_14-22-53

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27666 |
| Configuration ID: | 14346 | Run Name: | Run\_3-21\_14-22-53 |
| Draft Run: | N | State: |  |
| Duration: | 147 | Status: | Passed |
| Exec Date: | 21/03/17 | Target Cycle: |  |
| Exec Time: | 14:25:20 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27666]Finger Oximeter - General Screen | Test: Test Name: | Finger Oximeter - General Screen |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27665 - Run\_3-21\_13-50-28

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27665 |
| Configuration ID: | 14414 | Run Name: | Run\_3-21\_13-50-28 |
| Draft Run: | N | State: |  |
| Duration: | 947 | Status: | Passed |
| Exec Date: | 21/03/17 | Target Cycle: |  |
| Exec Time: | 14:06:15 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27665]Admin Configuration Screen | Test: Test Name: | Admin Configuration Screen |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27664 - Fast\_Run\_3-21\_13-21-19

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 27664 |
| Configuration ID: | 14395 | Run Name: | Fast\_Run\_3-21\_13-21-19 |
| Draft Run: | N | State: |  |
| Duration: | 80 | Status: | Passed |
| Exec Date: | 09/05/17 | Target Cycle: |  |
| Exec Time: | 11:31:09 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27664]Weight Threshold | Test: Test Name: | Weight Threshold |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27663 - Fast\_Run\_3-21\_13-21-16

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 27663 |
| Configuration ID: | 14384 | Run Name: | Fast\_Run\_3-21\_13-21-16 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Failed |
| Exec Date: | 21/03/17 | Target Cycle: |  |
| Exec Time: | 13:21:16 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27663]Threshold Input Value Validation | Test: Test Name: | Threshold Input Value Validation |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27662 - Fast\_Run\_3-21\_13-21-12

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 27662 |
| Configuration ID: | 14385 | Run Name: | Fast\_Run\_3-21\_13-21-12 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Failed |
| Exec Date: | 21/03/17 | Target Cycle: |  |
| Exec Time: | 13:21:12 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27662]Temperature Threshold | Test: Test Name: | Temperature Threshold |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27661 - Fast\_Run\_3-21\_13-21-9

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 27661 |
| Configuration ID: | 14386 | Run Name: | Fast\_Run\_3-21\_13-21-9 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Failed |
| Exec Date: | 21/03/17 | Target Cycle: |  |
| Exec Time: | 13:21:09 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27661]Pulse Rate Threshold | Test: Test Name: | Pulse Rate Threshold |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27660 - Fast\_Run\_3-21\_13-21-7

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 27660 |
| Configuration ID: | 14387 | Run Name: | Fast\_Run\_3-21\_13-21-7 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Failed |
| Exec Date: | 21/03/17 | Target Cycle: |  |
| Exec Time: | 13:21:07 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27660]Oxygen Saturation Threshold | Test: Test Name: | Oxygen Saturation Threshold |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27659 - Fast\_Run\_3-21\_13-21-4

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 27659 |
| Configuration ID: | 14396 | Run Name: | Fast\_Run\_3-21\_13-21-4 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Failed |
| Exec Date: | 21/03/17 | Target Cycle: |  |
| Exec Time: | 13:21:04 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27659]Blood Pressure Threshold | Test: Test Name: | Blood Pressure Threshold |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27656 - Run\_3-21\_11-51-5

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27656 |
| Configuration ID: | 14344 | Run Name: | Run\_3-21\_11-51-5 |
| Draft Run: | N | State: |  |
| Duration: | 6811 | Status: | Passed |
| Exec Date: | 21/03/17 | Target Cycle: |  |
| Exec Time: | 13:44:37 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27656]Logon - MediPi Session Lock | Test: Test Name: | Logon - MediPi Session Lock |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27655 - Run\_3-21\_11-49-29

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27655 |
| Configuration ID: | 14336 | Run Name: | Run\_3-21\_11-49-29 |
| Draft Run: | N | State: |  |
| Duration: | 83 | Status: | Passed |
| Exec Date: | 21/03/17 | Target Cycle: |  |
| Exec Time: | 11:50:53 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27655]Logon - General and Patient Logon | Test: Test Name: | Logon - General and Patient Logon |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27653 - Run\_3-21\_11-39-38

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27653 |
| Configuration ID: | 14361 | Run Name: | Run\_3-21\_11-39-38 |
| Draft Run: | N | State: |  |
| Duration: | 470 | Status: | Passed |
| Exec Date: | 21/03/17 | Target Cycle: |  |
| Exec Time: | 11:47:30 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27653]Homepage - General Screen | Test: Test Name: | Homepage - General Screen |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27652 - Run\_3-21\_11-33-9

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27652 |
| Configuration ID: | 14377 | Run Name: | Run\_3-21\_11-33-9 |
| Draft Run: | N | State: |  |
| Duration: | 48 | Status: | Passed |
| Exec Date: | 21/03/17 | Target Cycle: |  |
| Exec Time: | 11:33:58 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27652]Patients Caseload View | Test: Test Name: | Patients Caseload View |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27651 - Run\_3-21\_11-25-43

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27651 |
| Configuration ID: | 14379 | Run Name: | Run\_3-21\_11-25-43 |
| Draft Run: | N | State: |  |
| Duration: | 408 | Status: | Passed |
| Exec Date: | 21/03/17 | Target Cycle: |  |
| Exec Time: | 11:32:33 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27651]Patient Detail View | Test: Test Name: | Patient Detail View |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27650 - Run\_3-21\_11-22-48

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27650 |
| Configuration ID: | 14375 | Run Name: | Run\_3-21\_11-22-48 |
| Draft Run: | N | State: |  |
| Duration: | 84 | Status: | Failed |
| Exec Date: | 21/03/17 | Target Cycle: |  |
| Exec Time: | 11:24:13 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27650]Browser Testing - Internet Explorer | Test: Test Name: | Browser Testing - Internet Explorer |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27649 - Run\_3-21\_11-16-32

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27649 |
| Configuration ID: | 14376 | Run Name: | Run\_3-21\_11-16-32 |
| Draft Run: | N | State: |  |
| Duration: | 359 | Status: | Passed |
| Exec Date: | 21/03/17 | Target Cycle: |  |
| Exec Time: | 11:22:32 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27649]Homepage - General Screen | Test: Test Name: | Homepage - General Screen |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27648 - Run\_3-21\_11-10-1

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27648 |
| Configuration ID: | 14473 | Run Name: | Run\_3-21\_11-10-1 |
| Draft Run: | N | State: |  |
| Duration: | 167 | Status: | Passed |
| Exec Date: | 21/03/17 | Target Cycle: |  |
| Exec Time: | 11:12:49 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27648]Diabetes Questionnaire Results | Test: Test Name: | Diabetes Questionnaire Results |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27647 - Run\_3-21\_11-7-55

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27647 |
| Configuration ID: | 14472 | Run Name: | Run\_3-21\_11-7-55 |
| Draft Run: | N | State: |  |
| Duration: | 122 | Status: | Passed |
| Exec Date: | 21/03/17 | Target Cycle: |  |
| Exec Time: | 11:09:58 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27647]COPD Questionnaire Results | Test: Test Name: | COPD Questionnaire Results |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27646 - Run\_3-21\_11-2-59

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27646 |
| Configuration ID: | 14471 | Run Name: | Run\_3-21\_11-2-59 |
| Draft Run: | N | State: |  |
| Duration: | 289 | Status: | Passed |
| Exec Date: | 21/03/17 | Target Cycle: |  |
| Exec Time: | 11:07:50 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27646]Heart Questionnaire Results | Test: Test Name: | Heart Questionnaire Results |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27644 - Run\_3-21\_10-56-19

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27644 |
| Configuration ID: | 14469 | Run Name: | Run\_3-21\_10-56-19 |
| Draft Run: | N | State: |  |
| Duration: | 166 | Status: | Passed |
| Exec Date: | 21/03/17 | Target Cycle: |  |
| Exec Time: | 10:59:06 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.15 (Final fixes and Full Regression) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27644]Logon & Logoff | Test: Test Name: | Logon & Logoff |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27105 - Run\_1-27\_14-22-3

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27105 |
| Configuration ID: | 14416 | Run Name: | Run\_1-27\_14-22-3 |
| Draft Run: | N | State: |  |
| Duration: | 23 | Status: | Passed |
| Exec Date: | 27/01/17 | Target Cycle: |  |
| Exec Time: | 14:22:27 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.14 PILOT-20170127-1 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27105]Data transmitted from multiple MediPi Devices | Test: Test Name: | Data transmitted from multiple MediPi Devices |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27104 - Run\_1-27\_11-34-28

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27104 |
| Configuration ID: | 14418 | Run Name: | Run\_1-27\_11-34-28 |
| Draft Run: | N | State: |  |
| Duration: | 243 | Status: | Failed |
| Exec Date: | 27/01/17 | Target Cycle: |  |
| Exec Time: | 11:38:33 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.14 PILOT-20170127-1 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27104]Finger Oximeter - Time Synchronisation | Test: Test Name: | Finger Oximeter - Time Synchronisation |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27103 - Run\_1-27\_11-14-6

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27103 |
| Configuration ID: | 14420 | Run Name: | Run\_1-27\_11-14-6 |
| Draft Run: | N | State: |  |
| Duration: | 124 | Status: | Passed |
| Exec Date: | 27/01/17 | Target Cycle: |  |
| Exec Time: | 11:16:11 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.14 PILOT-20170127-1 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27103]Blood Pressure Monitor - Time Synchronisation | Test: Test Name: | Blood Pressure Monitor - Time Synchronisation |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27102 - Fast\_Run\_1-26\_12-42-49

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 27102 |
| Configuration ID: | 14397 | Run Name: | Fast\_Run\_1-26\_12-42-49 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 26/01/17 | Target Cycle: |  |
| Exec Time: | 12:42:49 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi v1.0.14 PILOT - 20170125 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27102]COPD Questionnaire Alerts | Test: Test Name: | COPD Questionnaire Alerts |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27101 - Run\_1-26\_12-42-12

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27101 |
| Configuration ID: | 14398 | Run Name: | Run\_1-26\_12-42-12 |
| Draft Run: | N | State: |  |
| Duration: | 6 | Status: | Passed |
| Exec Date: | 26/01/17 | Target Cycle: |  |
| Exec Time: | 12:42:18 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.14 PILOT - 20170125 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27101]Diabetes Questionnaire Alerts | Test: Test Name: | Diabetes Questionnaire Alerts |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27100 - Run\_1-26\_12-42-1

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27100 |
| Configuration ID: | 14405 | Run Name: | Run\_1-26\_12-42-1 |
| Draft Run: | N | State: |  |
| Duration: | 7 | Status: | Passed |
| Exec Date: | 26/01/17 | Target Cycle: |  |
| Exec Time: | 12:42:09 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.14 PILOT - 20170125 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27100]Diabetes Questionnaire - Transmit | Test: Test Name: | Diabetes Questionnaire - Transmit |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27099 - Run\_1-26\_12-41-47

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27099 |
| Configuration ID: | 14402 | Run Name: | Run\_1-26\_12-41-47 |
| Draft Run: | N | State: |  |
| Duration: | 11 | Status: | Passed |
| Exec Date: | 26/01/17 | Target Cycle: |  |
| Exec Time: | 12:41:58 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.14 PILOT - 20170125 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27099]Diabetes Questionnaire | Test: Test Name: | Diabetes Questionnaire |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27098 - Run\_1-26\_12-11-47

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27098 |
| Configuration ID: | 14371 | Run Name: | Run\_1-26\_12-11-47 |
| Draft Run: | N | State: |  |
| Duration: | 6 | Status: | Passed |
| Exec Date: | 26/01/17 | Target Cycle: |  |
| Exec Time: | 12:11:53 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.14 PILOT - 20170125 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27098]Heart Questionnaire - Transmit | Test: Test Name: | Heart Questionnaire - Transmit |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27097 - Run\_1-26\_12-11-29

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27097 |
| Configuration ID: | 14349 | Run Name: | Run\_1-26\_12-11-29 |
| Draft Run: | N | State: |  |
| Duration: | 13 | Status: | Passed |
| Exec Date: | 26/01/17 | Target Cycle: |  |
| Exec Time: | 12:11:43 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.14 PILOT - 20170125 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27097]Heart Failure Questionnaire | Test: Test Name: | Heart Failure Questionnaire |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27096 - Run\_1-26\_12-10-13

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27096 |
| Configuration ID: | 14388 | Run Name: | Run\_1-26\_12-10-13 |
| Draft Run: | N | State: |  |
| Duration: | 22 | Status: | Passed |
| Exec Date: | 26/01/17 | Target Cycle: |  |
| Exec Time: | 12:10:36 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.14 PILOT - 20170125 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27096]Heart Questionnaire - Alerts | Test: Test Name: | Heart Questionnaire - Alerts |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27095 - Fast\_Run\_1-26\_12-2-6

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 27095 |
| Configuration ID: | 14397 | Run Name: | Fast\_Run\_1-26\_12-2-6 |
| Draft Run: | N | State: |  |
| Duration: | 10 | Status: | Not Completed |
| Exec Date: | 26/01/17 | Target Cycle: |  |
| Exec Time: | 12:02:06 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.14 PILOT - 20170125 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27095]COPD Questionnaire Alerts | Test: Test Name: | COPD Questionnaire Alerts |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27094 - Run\_1-26\_11-21-56

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27094 |
| Configuration ID: | 14397 | Run Name: | Run\_1-26\_11-21-56 |
| Draft Run: | N | State: |  |
| Duration: | 2399 | Status: | Passed |
| Exec Date: | 26/01/17 | Target Cycle: |  |
| Exec Time: | 12:01:55 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.14 PILOT - 20170125 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27094]COPD Questionnaire Alerts | Test: Test Name: | COPD Questionnaire Alerts |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27093 - Run\_1-26\_11-20-51

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27093 |
| Configuration ID: | 14404 | Run Name: | Run\_1-26\_11-20-51 |
| Draft Run: | N | State: |  |
| Duration: | 62 | Status: | Passed |
| Exec Date: | 26/01/17 | Target Cycle: |  |
| Exec Time: | 11:21:53 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.14 PILOT - 20170125 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27093]COPD Questionnaire - Transmit | Test: Test Name: | COPD Questionnaire - Transmit |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27092 - Run\_1-26\_11-3-23

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27092 |
| Configuration ID: | 14403 | Run Name: | Run\_1-26\_11-3-23 |
| Draft Run: | N | State: |  |
| Duration: | 1039 | Status: | Passed |
| Exec Date: | 26/01/17 | Target Cycle: |  |
| Exec Time: | 11:20:43 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.14 PILOT - 20170125 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27092]COPD Questionnaire | Test: Test Name: | COPD Questionnaire |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27091 - Run\_1-26\_11-1-45

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27091 |
| Configuration ID: | 14389 | Run Name: | Run\_1-26\_11-1-45 |
| Draft Run: | N | State: |  |
| Duration: | 14 | Status: | Passed |
| Exec Date: | 26/01/17 | Target Cycle: |  |
| Exec Time: | 11:01:59 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.14 PILOT - 20170125 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27091]Caseload View - Question Mark | Test: Test Name: | Caseload View - Question Mark |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27090 - Run\_1-26\_11-0-54

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27090 |
| Configuration ID: | 14394 | Run Name: | Run\_1-26\_11-0-54 |
| Draft Run: | N | State: |  |
| Duration: | 33 | Status: | Passed |
| Exec Date: | 26/01/17 | Target Cycle: |  |
| Exec Time: | 11:01:38 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.14 PILOT - 20170125 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27090]Caseload View - Question Mark Reset | Test: Test Name: | Caseload View - Question Mark Reset |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27089 - Run\_1-26\_11-0-36

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27089 |
| Configuration ID: | 14406 | Run Name: | Run\_1-26\_11-0-36 |
| Draft Run: | N | State: |  |
| Duration: | 9 | Status: | Passed |
| Exec Date: | 26/01/17 | Target Cycle: |  |
| Exec Time: | 11:00:45 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.14 PILOT - 20170125 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27089]Scheduler Configuration | Test: Test Name: | Scheduler Configuration |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27087 - Run\_1-25\_15-35-49

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27087 |
| Configuration ID: | 14393 | Run Name: | Run\_1-25\_15-35-49 |
| Draft Run: | N | State: |  |
| Duration: | 768 | Status: | Passed |
| Exec Date: | 25/01/17 | Target Cycle: |  |
| Exec Time: | 15:48:38 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.14 PILOT - 20170125 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27087]Caseload View - Sad Face | Test: Test Name: | Caseload View - Sad Face |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27086 - Run\_1-25\_15-30-38

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27086 |
| Configuration ID: | 14378 | Run Name: | Run\_1-25\_15-30-38 |
| Draft Run: | N | State: |  |
| Duration: | 306 | Status: | Passed |
| Exec Date: | 25/01/17 | Target Cycle: |  |
| Exec Time: | 15:35:44 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.14 PILOT - 20170125 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27086]Caseload View - Happy Face | Test: Test Name: | Caseload View - Happy Face |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27085 - Run\_1-25\_15-27-58

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27085 |
| Configuration ID: | 14391 | Run Name: | Run\_1-25\_15-27-58 |
| Draft Run: | N | State: |  |
| Duration: | 97 | Status: | Passed |
| Exec Date: | 25/01/17 | Target Cycle: |  |
| Exec Time: | 15:29:37 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.14 PILOT - 20170125 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27085]Caseload View - Exclamation Mark | Test: Test Name: | Caseload View - Exclamation Mark |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27084 - Run\_1-25\_14-17-43

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27084 |
| Configuration ID: | 14378 | Run Name: | Run\_1-25\_14-17-43 |
| Draft Run: | N | State: |  |
| Duration: | 133 | Status: | Passed |
| Exec Date: | 25/01/17 | Target Cycle: |  |
| Exec Time: | 14:19:57 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.14 PILOT - 20170125 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27084]Caseload View - Happy Face | Test: Test Name: | Caseload View - Happy Face |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27083 - Run\_1-25\_13-40-51

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27083 |
| Configuration ID: | 14396 | Run Name: | Run\_1-25\_13-40-51 |
| Draft Run: | N | State: |  |
| Duration: | 1173 | Status: | Failed |
| Exec Date: | 26/01/17 | Target Cycle: |  |
| Exec Time: | 12:03:41 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.14 PILOT - 20170125 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27083]Blood Pressure Threshold | Test: Test Name: | Blood Pressure Threshold |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27081 - Run\_1-25\_11-35-16

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27081 |
| Configuration ID: | 14395 | Run Name: | Run\_1-25\_11-35-16 |
| Draft Run: | N | State: |  |
| Duration: | 2175 | Status: | Failed |
| Exec Date: | 25/01/17 | Target Cycle: |  |
| Exec Time: | 12:11:32 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.14 PILOT - 20170125 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27081]Weight Threshold | Test: Test Name: | Weight Threshold |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27080 - Run\_1-25\_11-29-13

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27080 |
| Configuration ID: | 14385 | Run Name: | Run\_1-25\_11-29-13 |
| Draft Run: | N | State: |  |
| Duration: | 352 | Status: | Passed |
| Exec Date: | 25/01/17 | Target Cycle: |  |
| Exec Time: | 11:35:07 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.14 PILOT - 20170125 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27080]Temperature Threshold | Test: Test Name: | Temperature Threshold |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27079 - Run\_1-25\_11-7-47

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27079 |
| Configuration ID: | 14386 | Run Name: | Run\_1-25\_11-7-47 |
| Draft Run: | N | State: |  |
| Duration: | 293 | Status: | Passed |
| Exec Date: | 25/01/17 | Target Cycle: |  |
| Exec Time: | 11:12:41 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.14 PILOT - 20170125 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27079]Pulse Rate Threshold | Test: Test Name: | Pulse Rate Threshold |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27078 - Run\_1-25\_10-38-55

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27078 |
| Configuration ID: | 14387 | Run Name: | Run\_1-25\_10-38-55 |
| Draft Run: | N | State: |  |
| Duration: | 1650 | Status: | Passed |
| Exec Date: | 25/01/17 | Target Cycle: |  |
| Exec Time: | 11:06:26 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.14 PILOT - 20170125 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27078]Oxygen Saturation Threshold | Test: Test Name: | Oxygen Saturation Threshold |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27077 - Run\_1-25\_10-37-16

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27077 |
| Configuration ID: | 14417 | Run Name: | Run\_1-25\_10-37-16 |
| Draft Run: | N | State: |  |
| Duration: | 16 | Status: | Passed |
| Exec Date: | 25/01/17 | Target Cycle: |  |
| Exec Time: | 10:37:33 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.14 PILOT - 20170125 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27077]Loss of WIFI Connection | Test: Test Name: | Loss of WIFI Connection |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27076 - Run\_1-24\_15-29-20

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27076 |
| Configuration ID: | 14415 | Run Name: | Run\_1-24\_15-29-20 |
| Draft Run: | N | State: |  |
| Duration: | 1229 | Status: | Passed |
| Exec Date: | 24/01/17 | Target Cycle: |  |
| Exec Time: | 15:49:50 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.14 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27076]MediPi Time Synchronisation | Test: Test Name: | MediPi Time Synchronisation |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27075 - Fast\_Run\_1-24\_15-29-20

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 27075 |
| Configuration ID: | 14377 | Run Name: | Fast\_Run\_1-24\_15-29-20 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Passed |
| Exec Date: | 24/01/17 | Target Cycle: |  |
| Exec Time: | 15:29:20 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi v1.0.14 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27075]Patients Caseload View | Test: Test Name: | Patients Caseload View |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27073 - Run\_1-24\_11-46-10

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27073 |
| Configuration ID: | 14336 | Run Name: | Run\_1-24\_11-46-10 |
| Draft Run: | N | State: |  |
| Duration: | 326 | Status: | Passed |
| Exec Date: | 24/01/17 | Target Cycle: |  |
| Exec Time: | 11:51:37 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.14 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27073]Logon - General and Patient Logon | Test: Test Name: | Logon - General and Patient Logon |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27071 - Run\_1-24\_11-35-30

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27071 |
| Configuration ID: | 14414 | Run Name: | Run\_1-24\_11-35-30 |
| Draft Run: | N | State: |  |
| Duration: | 339 | Status: | Failed |
| Exec Date: | 24/01/17 | Target Cycle: |  |
| Exec Time: | 11:41:10 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.14 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27071]Admin Configuration Screen | Test: Test Name: | Admin Configuration Screen |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27069 - Run\_1-24\_9-9-45

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27069 |
| Configuration ID: | 14377 | Run Name: | Run\_1-24\_9-9-45 |
| Draft Run: | N | State: |  |
| Duration: | 113 | Status: | Not Completed |
| Exec Date: | 24/01/17 | Target Cycle: |  |
| Exec Time: | 09:11:39 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.14 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27069]Patients Caseload View | Test: Test Name: | Patients Caseload View |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27068 - Run\_1-24\_9-7-37

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27068 |
| Configuration ID: | 14379 | Run Name: | Run\_1-24\_9-7-37 |
| Draft Run: | N | State: |  |
| Duration: | 118 | Status: | Passed |
| Exec Date: | 24/01/17 | Target Cycle: |  |
| Exec Time: | 09:09:36 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.14 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27068]Patient Detail View | Test: Test Name: | Patient Detail View |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27067 - Run\_1-24\_9-6-54

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27067 |
| Configuration ID: | 14376 | Run Name: | Run\_1-24\_9-6-54 |
| Draft Run: | N | State: |  |
| Duration: | 38 | Status: | Passed |
| Exec Date: | 24/01/17 | Target Cycle: |  |
| Exec Time: | 09:07:34 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.14 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27067]Homepage - General Screen | Test: Test Name: | Homepage - General Screen |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27027 - Run\_12-21\_10-54-33

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27027 |
| Configuration ID: | 14375 | Run Name: | Run\_12-21\_10-54-33 |
| Draft Run: | N | State: |  |
| Duration: | 839 | Status: | Failed |
| Exec Date: | 21/12/16 | Target Cycle: |  |
| Exec Time: | 11:08:33 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27027]Browser Testing - Internet Explorer | Test: Test Name: | Browser Testing - Internet Explorer |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27022 - Fast\_Run\_12-21\_10-30-23

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 27022 |
| Configuration ID: | 14394 | Run Name: | Fast\_Run\_12-21\_10-30-23 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Blocked |
| Exec Date: | 21/12/16 | Target Cycle: |  |
| Exec Time: | 10:30:23 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27022]Caseload View - Question Mark Reset | Test: Test Name: | Caseload View - Question Mark Reset |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27021 - Fast\_Run\_12-21\_10-30-20

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 27021 |
| Configuration ID: | 14389 | Run Name: | Fast\_Run\_12-21\_10-30-20 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Blocked |
| Exec Date: | 21/12/16 | Target Cycle: |  |
| Exec Time: | 10:30:20 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27021]Caseload View - Question Mark | Test: Test Name: | Caseload View - Question Mark |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27020 - Fast\_Run\_12-21\_10-23-56

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 27020 |
| Configuration ID: | 14393 | Run Name: | Fast\_Run\_12-21\_10-23-56 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Blocked |
| Exec Date: | 21/12/16 | Target Cycle: |  |
| Exec Time: | 10:23:56 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27020]Caseload View - Sad Face | Test: Test Name: | Caseload View - Sad Face |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27019 - Run\_12-21\_10-19-2

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27019 |
| Configuration ID: | 14393 | Run Name: | Run\_12-21\_10-19-2 |
| Draft Run: | N | State: |  |
| Duration: | 280 | Status: | Not Completed |
| Exec Date: | 21/12/16 | Target Cycle: |  |
| Exec Time: | 10:23:43 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27019]Caseload View - Sad Face | Test: Test Name: | Caseload View - Sad Face |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27018 - Fast\_Run\_12-21\_10-18-35

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 27018 |
| Configuration ID: | 14378 | Run Name: | Fast\_Run\_12-21\_10-18-35 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Blocked |
| Exec Date: | 21/12/16 | Target Cycle: |  |
| Exec Time: | 10:18:35 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27018]Caseload View - Happy Face | Test: Test Name: | Caseload View - Happy Face |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27017 - Run\_12-21\_9-53-42

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27017 |
| Configuration ID: | 14378 | Run Name: | Run\_12-21\_9-53-42 |
| Draft Run: | N | State: |  |
| Duration: | 1474 | Status: | Not Completed |
| Exec Date: | 21/12/16 | Target Cycle: |  |
| Exec Time: | 10:18:16 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27017]Caseload View - Happy Face | Test: Test Name: | Caseload View - Happy Face |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27013 - Fast\_Run\_12-21\_9-46-41

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 27013 |
| Configuration ID: | 14391 | Run Name: | Fast\_Run\_12-21\_9-46-41 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Blocked |
| Exec Date: | 21/12/16 | Target Cycle: |  |
| Exec Time: | 09:46:41 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27013]Caseload View - Exclamation Mark | Test: Test Name: | Caseload View - Exclamation Mark |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27012 - Run\_12-21\_9-40-48

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27012 |
| Configuration ID: | 14391 | Run Name: | Run\_12-21\_9-40-48 |
| Draft Run: | N | State: |  |
| Duration: | 277 | Status: | Not Completed |
| Exec Date: | 21/12/16 | Target Cycle: |  |
| Exec Time: | 09:45:26 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27012]Caseload View - Exclamation Mark | Test: Test Name: | Caseload View - Exclamation Mark |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27011 - Run\_12-20\_13-4-41

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27011 |
| Configuration ID: | 14395 | Run Name: | Run\_12-20\_13-4-41 |
| Draft Run: | N | State: |  |
| Duration: | 216 | Status: | Failed |
| Exec Date: | 21/12/16 | Target Cycle: |  |
| Exec Time: | 11:11:04 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27011]Weight Threshold | Test: Test Name: | Weight Threshold |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27010 - Run\_12-20\_13-2-19

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27010 |
| Configuration ID: | 14387 | Run Name: | Run\_12-20\_13-2-19 |
| Draft Run: | N | State: |  |
| Duration: | 786 | Status: | Failed |
| Exec Date: | 21/12/16 | Target Cycle: |  |
| Exec Time: | 09:23:25 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27010]Oxygen Saturation Threshold | Test: Test Name: | Oxygen Saturation Threshold |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 27000 - Run\_12-19\_15-30-53

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 27000 |
| Configuration ID: | 14396 | Run Name: | Run\_12-19\_15-30-53 |
| Draft Run: | N | State: |  |
| Duration: | 2038 | Status: | Blocked |
| Exec Date: | 21/12/16 | Target Cycle: |  |
| Exec Time: | 09:39:10 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [27000]Blood Pressure Threshold | Test: Test Name: | Blood Pressure Threshold |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26998 - Run\_12-19\_15-23-15

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26998 |
| Configuration ID: | 14385 | Run Name: | Run\_12-19\_15-23-15 |
| Draft Run: | N | State: |  |
| Duration: | 2055 | Status: | Failed |
| Exec Date: | 19/12/16 | Target Cycle: |  |
| Exec Time: | 15:30:20 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26998]Temperature Threshold | Test: Test Name: | Temperature Threshold |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26995 - Run\_12-19\_15-12-6

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26995 |
| Configuration ID: | 14386 | Run Name: | Run\_12-19\_15-12-6 |
| Draft Run: | N | State: |  |
| Duration: | 1680 | Status: | Failed |
| Exec Date: | 21/12/16 | Target Cycle: |  |
| Exec Time: | 09:11:15 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26995]Pulse Rate Threshold | Test: Test Name: | Pulse Rate Threshold |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26980 - Run\_12-19\_13-40-19

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26980 |
| Configuration ID: | 14384 | Run Name: | Run\_12-19\_13-40-19 |
| Draft Run: | N | State: |  |
| Duration: | 222 | Status: | Failed |
| Exec Date: | 19/12/16 | Target Cycle: |  |
| Exec Time: | 13:44:02 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26980]Threshold Input Value Validation | Test: Test Name: | Threshold Input Value Validation |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26979 - Run\_12-19\_12-13-21

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26979 |
| Configuration ID: | 14365 | Run Name: | Run\_12-19\_12-13-21 |
| Draft Run: | N | State: |  |
| Duration: | 216 | Status: | Failed |
| Exec Date: | 19/12/16 | Target Cycle: |  |
| Exec Time: | 12:16:58 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26979]MediPi Patient Banner - Patient Name | Test: Test Name: | MediPi Patient Banner - Patient Name |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26977 - Run\_12-19\_12-12-13

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26977 |
| Configuration ID: | 14364 | Run Name: | Run\_12-19\_12-12-13 |
| Draft Run: | N | State: |  |
| Duration: | 32 | Status: | Passed |
| Exec Date: | 19/12/16 | Target Cycle: |  |
| Exec Time: | 12:12:46 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26977]Fatal Error Screen | Test: Test Name: | Fatal Error Screen |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26976 - Run\_12-19\_12-11-19

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26976 |
| Configuration ID: | 14363 | Run Name: | Run\_12-19\_12-11-19 |
| Draft Run: | N | State: |  |
| Duration: | 42 | Status: | Passed |
| Exec Date: | 19/12/16 | Target Cycle: |  |
| Exec Time: | 12:12:02 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26976]Multiple Connected Devices | Test: Test Name: | Multiple Connected Devices |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26975 - Run\_12-19\_12-9-14

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26975 |
| Configuration ID: | 14377 | Run Name: | Run\_12-19\_12-9-14 |
| Draft Run: | N | State: |  |
| Duration: | 113 | Status: | Passed |
| Exec Date: | 19/12/16 | Target Cycle: |  |
| Exec Time: | 12:11:08 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26975]Patients Caseload View | Test: Test Name: | Patients Caseload View |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26974 - Run\_12-19\_12-3-30

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26974 |
| Configuration ID: | 14379 | Run Name: | Run\_12-19\_12-3-30 |
| Draft Run: | N | State: |  |
| Duration: | 275 | Status: | Passed |
| Exec Date: | 19/12/16 | Target Cycle: |  |
| Exec Time: | 12:08:06 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26974]Patient Detail View | Test: Test Name: | Patient Detail View |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26973 - Run\_12-19\_12-2-34

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26973 |
| Configuration ID: | 14376 | Run Name: | Run\_12-19\_12-2-34 |
| Draft Run: | N | State: |  |
| Duration: | 51 | Status: | Passed |
| Exec Date: | 19/12/16 | Target Cycle: |  |
| Exec Time: | 12:03:26 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26973]Homepage - General Screen | Test: Test Name: | Homepage - General Screen |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26972 - Run\_12-19\_11-59-32

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26972 |
| Configuration ID: | 14360 | Run Name: | Run\_12-19\_11-59-32 |
| Draft Run: | N | State: |  |
| Duration: | 144 | Status: | Passed |
| Exec Date: | 19/12/16 | Target Cycle: |  |
| Exec Time: | 12:01:57 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26972]Scheduler - Transmit Summary Screen | Test: Test Name: | Scheduler - Transmit Summary Screen |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26971 - Run\_12-19\_11-50-42

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26971 |
| Configuration ID: | 14373 | Run Name: | Run\_12-19\_11-50-42 |
| Draft Run: | N | State: |  |
| Duration: | 524 | Status: | Passed |
| Exec Date: | 19/12/16 | Target Cycle: |  |
| Exec Time: | 11:59:27 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26971]Scheduler - Transmit (All Data) | Test: Test Name: | Scheduler - Transmit (All Data) |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26970 - Run\_12-19\_11-39-33

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26970 |
| Configuration ID: | 14374 | Run Name: | Run\_12-19\_11-39-33 |
| Draft Run: | N | State: |  |
| Duration: | 662 | Status: | Passed |
| Exec Date: | 19/12/16 | Target Cycle: |  |
| Exec Time: | 11:50:36 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26970]Scheduler - Transmit (Individual) | Test: Test Name: | Scheduler - Transmit (Individual) |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26915 - Run\_12-15\_15-6-46

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26915 |
| Configuration ID: | 14367 | Run Name: | Run\_12-15\_15-6-46 |
| Draft Run: | N | State: |  |
| Duration: | 2856 | Status: | Passed |
| Exec Date: | 19/12/16 | Target Cycle: |  |
| Exec Time: | 11:39:20 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26915]Transmitter - Transmit (All Data) | Test: Test Name: | Transmitter - Transmit (All Data) |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26910 - Run\_12-15\_13-59-45

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26910 |
| Configuration ID: | 14397 | Run Name: | Run\_12-15\_13-59-45 |
| Draft Run: | N | State: |  |
| Duration: | 70 | Status: | Failed |
| Exec Date: | 15/12/16 | Target Cycle: |  |
| Exec Time: | 14:00:55 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26910]COPD Questionnaire Alerts | Test: Test Name: | COPD Questionnaire Alerts |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26908 - Run\_12-15\_13-56-13

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26908 |
| Configuration ID: | 14398 | Run Name: | Run\_12-15\_13-56-13 |
| Draft Run: | N | State: |  |
| Duration: | 135 | Status: | Failed |
| Exec Date: | 15/12/16 | Target Cycle: |  |
| Exec Time: | 13:58:29 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26908]Diabetes Questionnaire Alerts | Test: Test Name: | Diabetes Questionnaire Alerts |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26907 - Run\_12-15\_13-48-27

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26907 |
| Configuration ID: | 14388 | Run Name: | Run\_12-15\_13-48-27 |
| Draft Run: | N | State: |  |
| Duration: | 98 | Status: | Failed |
| Exec Date: | 15/12/16 | Target Cycle: |  |
| Exec Time: | 13:50:07 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26907]Heart Questionnaire - Alerts | Test: Test Name: | Heart Questionnaire - Alerts |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26897 - Run\_12-15\_11-55-17

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26897 |
| Configuration ID: | 14350 | Run Name: | Run\_12-15\_11-55-17 |
| Draft Run: | N | State: |  |
| Duration: | 928 | Status: | Passed |
| Exec Date: | 15/12/16 | Target Cycle: |  |
| Exec Time: | 12:10:46 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26897]Scheduler - Schedule Screens | Test: Test Name: | Scheduler - Schedule Screens |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26895 - Run\_12-15\_11-50-38

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26895 |
| Configuration ID: | 14352 | Run Name: | Run\_12-15\_11-50-38 |
| Draft Run: | N | State: |  |
| Duration: | 256 | Status: | Passed |
| Exec Date: | 15/12/16 | Target Cycle: |  |
| Exec Time: | 11:54:55 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26895]Transmitter - General Screen | Test: Test Name: | Transmitter - General Screen |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26893 - Run\_12-15\_11-48-51

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26893 |
| Configuration ID: | 14366 | Run Name: | Run\_12-15\_11-48-51 |
| Draft Run: | N | State: |  |
| Duration: | 24 | Status: | Passed |
| Exec Date: | 15/12/16 | Target Cycle: |  |
| Exec Time: | 11:49:16 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26893]Clinician Messages - Polling Warning | Test: Test Name: | Clinician Messages - Polling Warning |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26887 - Run\_12-15\_11-20-1

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26887 |
| Configuration ID: | 14351 | Run Name: | Run\_12-15\_11-20-1 |
| Draft Run: | N | State: |  |
| Duration: | 261 | Status: | Passed |
| Exec Date: | 15/12/16 | Target Cycle: |  |
| Exec Time: | 11:24:23 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26887]Clinician Messages - General Screen | Test: Test Name: | Clinician Messages - General Screen |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26884 - Run\_12-15\_11-17-59

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26884 |
| Configuration ID: | 14404 | Run Name: | Run\_12-15\_11-17-59 |
| Draft Run: | N | State: |  |
| Duration: | 20 | Status: | Passed |
| Exec Date: | 15/12/16 | Target Cycle: |  |
| Exec Time: | 11:18:20 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26884]COPD Questionnaire - Transmit | Test: Test Name: | COPD Questionnaire - Transmit |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26883 - Run\_12-15\_11-17-16

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26883 |
| Configuration ID: | 14405 | Run Name: | Run\_12-15\_11-17-16 |
| Draft Run: | N | State: |  |
| Duration: | 25 | Status: | Passed |
| Exec Date: | 15/12/16 | Target Cycle: |  |
| Exec Time: | 11:17:42 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26883]Diabetes Questionnaire - Transmit | Test: Test Name: | Diabetes Questionnaire - Transmit |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26882 - Run\_12-15\_11-14-42

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26882 |
| Configuration ID: | 14371 | Run Name: | Run\_12-15\_11-14-42 |
| Draft Run: | N | State: |  |
| Duration: | 108 | Status: | Passed |
| Exec Date: | 15/12/16 | Target Cycle: |  |
| Exec Time: | 11:16:31 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26882]Heart Questionnaire - Transmit | Test: Test Name: | Heart Questionnaire - Transmit |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26880 - Run\_12-15\_11-6-13

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26880 |
| Configuration ID: | 14402 | Run Name: | Run\_12-15\_11-6-13 |
| Draft Run: | N | State: |  |
| Duration: | 43 | Status: | Passed |
| Exec Date: | 15/12/16 | Target Cycle: |  |
| Exec Time: | 11:06:57 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26880]Diabetes Questionnaire | Test: Test Name: | Diabetes Questionnaire |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26879 - Run\_12-15\_11-3-42

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26879 |
| Configuration ID: | 14403 | Run Name: | Run\_12-15\_11-3-42 |
| Draft Run: | N | State: |  |
| Duration: | 115 | Status: | Passed |
| Exec Date: | 15/12/16 | Target Cycle: |  |
| Exec Time: | 11:05:38 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26879]COPD Questionnaire | Test: Test Name: | COPD Questionnaire |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26870 - Run\_12-15\_10-24-58

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26870 |
| Configuration ID: | 14349 | Run Name: | Run\_12-15\_10-24-58 |
| Draft Run: | N | State: |  |
| Duration: | 204 | Status: | Passed |
| Exec Date: | 15/12/16 | Target Cycle: |  |
| Exec Time: | 10:28:22 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26870]Heart Failure Questionnaire | Test: Test Name: | Heart Failure Questionnaire |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26868 - Run\_12-15\_10-14-6

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26868 |
| Configuration ID: | 14369 | Run Name: | Run\_12-15\_10-14-6 |
| Draft Run: | N | State: |  |
| Duration: | 436 | Status: | Failed |
| Exec Date: | 15/12/16 | Target Cycle: |  |
| Exec Time: | 10:22:08 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26868]Diagnostic Scale - Transmit | Test: Test Name: | Diagnostic Scale - Transmit |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26866 - Run\_12-15\_10-7-24

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26866 |
| Configuration ID: | 14358 | Run Name: | Run\_12-15\_10-7-24 |
| Draft Run: | N | State: |  |
| Duration: | 397 | Status: | Passed |
| Exec Date: | 15/12/16 | Target Cycle: |  |
| Exec Time: | 10:14:02 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26866]Diagnostic Scale - Record | Test: Test Name: | Diagnostic Scale - Record |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26865 - Run\_12-15\_10-5-32

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26865 |
| Configuration ID: | 14347 | Run Name: | Run\_12-15\_10-5-32 |
| Draft Run: | N | State: |  |
| Duration: | 107 | Status: | Passed |
| Exec Date: | 15/12/16 | Target Cycle: |  |
| Exec Time: | 10:07:20 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26865]Diagnostic Scale - General Screen | Test: Test Name: | Diagnostic Scale - General Screen |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26863 - Run\_12-15\_10-3-53

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26863 |
| Configuration ID: | 14370 | Run Name: | Run\_12-15\_10-3-53 |
| Draft Run: | N | State: |  |
| Duration: | 85 | Status: | Passed |
| Exec Date: | 15/12/16 | Target Cycle: |  |
| Exec Time: | 10:05:10 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26863]Blood Pressure - Transmit | Test: Test Name: | Blood Pressure - Transmit |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26861 - Run\_12-15\_9-56-40

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26861 |
| Configuration ID: | 14359 | Run Name: | Run\_12-15\_9-56-40 |
| Draft Run: | N | State: |  |
| Duration: | 429 | Status: | Passed |
| Exec Date: | 15/12/16 | Target Cycle: |  |
| Exec Time: | 10:03:50 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26861]Blood Pressure - Record | Test: Test Name: | Blood Pressure - Record |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26859 - Run\_12-15\_9-53-35

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26859 |
| Configuration ID: | 14348 | Run Name: | Run\_12-15\_9-53-35 |
| Draft Run: | N | State: |  |
| Duration: | 180 | Status: | Failed |
| Exec Date: | 15/12/16 | Target Cycle: |  |
| Exec Time: | 09:56:36 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26859]Blood Pressure - General Screen | Test: Test Name: | Blood Pressure - General Screen |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26858 - Run\_12-15\_9-51-53

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26858 |
| Configuration ID: | 14382 | Run Name: | Run\_12-15\_9-51-53 |
| Draft Run: | N | State: |  |
| Duration: | 107 | Status: | Passed |
| Exec Date: | 15/12/16 | Target Cycle: |  |
| Exec Time: | 09:53:29 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26858]Thermometer - Transmit | Test: Test Name: | Thermometer - Transmit |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26857 - Run\_12-15\_9-50-3

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26857 |
| Configuration ID: | 14381 | Run Name: | Run\_12-15\_9-50-3 |
| Draft Run: | N | State: |  |
| Duration: | 103 | Status: | Passed |
| Exec Date: | 15/12/16 | Target Cycle: |  |
| Exec Time: | 09:51:47 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26857]Thermometer - Record | Test: Test Name: | Thermometer - Record |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26856 - Run\_12-15\_9-47-28

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26856 |
| Configuration ID: | 14380 | Run Name: | Run\_12-15\_9-47-28 |
| Draft Run: | N | State: |  |
| Duration: | 149 | Status: | Passed |
| Exec Date: | 15/12/16 | Target Cycle: |  |
| Exec Time: | 09:49:58 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26856]Thermometer - General Screen | Test: Test Name: | Thermometer - General Screen |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26845 - Run\_12-14\_15-29-44

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26845 |
| Configuration ID: | 14362 | Run Name: | Run\_12-14\_15-29-44 |
| Draft Run: | N | State: |  |
| Duration: | 1678 | Status: | Failed |
| Exec Date: | 15/12/16 | Target Cycle: |  |
| Exec Time: | 09:47:13 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26845]Tile Configuration | Test: Test Name: | Tile Configuration |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26844 - Run\_12-14\_13-36-23

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26844 |
| Configuration ID: | 14344 | Run Name: | Run\_12-14\_13-36-23 |
| Draft Run: | N | State: |  |
| Duration: | 577 | Status: | Passed |
| Exec Date: | 14/12/16 | Target Cycle: |  |
| Exec Time: | 15:16:18 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26844]Logon - MediPi Session Lock | Test: Test Name: | Logon - MediPi Session Lock |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26843 - Run\_12-14\_13-33-47

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26843 |
| Configuration ID: | 14336 | Run Name: | Run\_12-14\_13-33-47 |
| Draft Run: | N | State: |  |
| Duration: | 124 | Status: | Passed |
| Exec Date: | 14/12/16 | Target Cycle: |  |
| Exec Time: | 13:35:51 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26843]Logon - General and Patient Logon | Test: Test Name: | Logon - General and Patient Logon |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26842 - Run\_12-14\_13-31-28

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26842 |
| Configuration ID: | 14361 | Run Name: | Run\_12-14\_13-31-28 |
| Draft Run: | N | State: |  |
| Duration: | 128 | Status: | Passed |
| Exec Date: | 14/12/16 | Target Cycle: |  |
| Exec Time: | 13:33:36 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26842]Homepage - General Screen | Test: Test Name: | Homepage - General Screen |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26841 - Run\_12-14\_13-15-34

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26841 |
| Configuration ID: | 14368 | Run Name: | Run\_12-14\_13-15-34 |
| Draft Run: | N | State: |  |
| Duration: | 1155 | Status: | Passed |
| Exec Date: | 15/12/16 | Target Cycle: |  |
| Exec Time: | 09:35:55 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26841]Finger Oximeter - Transmit | Test: Test Name: | Finger Oximeter - Transmit |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26840 - Run\_12-14\_13-12-28

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26840 |
| Configuration ID: | 14353 | Run Name: | Run\_12-14\_13-12-28 |
| Draft Run: | N | State: |  |
| Duration: | 137 | Status: | Passed |
| Exec Date: | 14/12/16 | Target Cycle: |  |
| Exec Time: | 13:14:45 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26840]Finger Oximeter - Record | Test: Test Name: | Finger Oximeter - Record |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26839 - Run\_12-14\_13-11-33

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26839 |
| Configuration ID: | 14346 | Run Name: | Run\_12-14\_13-11-33 |
| Draft Run: | N | State: |  |
| Duration: | 44 | Status: | Passed |
| Exec Date: | 14/12/16 | Target Cycle: |  |
| Exec Time: | 13:12:17 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi v1.0.12 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26839]Finger Oximeter - General Screen | Test: Test Name: | Finger Oximeter - General Screen |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26814 - Run\_12-12\_11-44-29

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26814 |
| Configuration ID: | 14380 | Run Name: | Run\_12-12\_11-44-29 |
| Draft Run: | N | State: |  |
| Duration: | 291 | Status: | Passed |
| Exec Date: | 12/12/16 | Target Cycle: |  |
| Exec Time: | 11:49:21 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi V1.0.11 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26814]Thermometer - General Screen | Test: Test Name: | Thermometer - General Screen |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26813 - Run\_12-12\_10-58-10

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26813 |
| Configuration ID: | 14370 | Run Name: | Run\_12-12\_10-58-10 |
| Draft Run: | N | State: |  |
| Duration: | 2762 | Status: | Passed |
| Exec Date: | 12/12/16 | Target Cycle: |  |
| Exec Time: | 11:44:13 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi V1.0.11 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26813]Blood Pressure - Transmit | Test: Test Name: | Blood Pressure - Transmit |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26812 - Run\_12-8\_14-12-42

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26812 |
| Configuration ID: | 14359 | Run Name: | Run\_12-8\_14-12-42 |
| Draft Run: | N | State: |  |
| Duration: | 2592 | Status: | Passed |
| Exec Date: | 12/12/16 | Target Cycle: |  |
| Exec Time: | 10:51:36 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi V1.0.11 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26812]Blood Pressure - Record | Test: Test Name: | Blood Pressure - Record |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26811 - Run\_12-8\_14-7-28

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26811 |
| Configuration ID: | 14348 | Run Name: | Run\_12-8\_14-7-28 |
| Draft Run: | N | State: |  |
| Duration: | 242 | Status: | Passed |
| Exec Date: | 08/12/16 | Target Cycle: |  |
| Exec Time: | 14:11:32 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi V1.0.11 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26811]Blood Pressure - General Screen | Test: Test Name: | Blood Pressure - General Screen |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26810 - Run\_12-8\_11-54-21

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26810 |
| Configuration ID: | 14368 | Run Name: | Run\_12-8\_11-54-21 |
| Draft Run: | N | State: |  |
| Duration: | 7465 | Status: | Passed |
| Exec Date: | 08/12/16 | Target Cycle: |  |
| Exec Time: | 13:58:47 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi V1.0.11 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26810]Finger Oximeter - Transmit | Test: Test Name: | Finger Oximeter - Transmit |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26809 - Run\_12-8\_11-35-17

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26809 |
| Configuration ID: | 14353 | Run Name: | Run\_12-8\_11-35-17 |
| Draft Run: | N | State: |  |
| Duration: | 1039 | Status: | Failed |
| Exec Date: | 08/12/16 | Target Cycle: |  |
| Exec Time: | 11:52:37 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi V1.0.11 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26809]Finger Oximeter - Record | Test: Test Name: | Finger Oximeter - Record |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26808 - Run\_12-8\_11-29-28

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26808 |
| Configuration ID: | 14346 | Run Name: | Run\_12-8\_11-29-28 |
| Draft Run: | N | State: |  |
| Duration: | 341 | Status: | Passed |
| Exec Date: | 08/12/16 | Target Cycle: |  |
| Exec Time: | 11:35:10 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi V1.0.11 (Bluetooth Connected Devices) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26808]Finger Oximeter - General Screen | Test: Test Name: | Finger Oximeter - General Screen |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26798 - Run\_10-28\_11-41-21

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26798 |
| Configuration ID: | 14369 | Run Name: | Run\_10-28\_11-41-21 |
| Draft Run: | N | State: |  |
| Duration: | 238 | Status: | Passed |
| Exec Date: | 28/10/16 | Target Cycle: |  |
| Exec Time: | 11:45:20 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_1.0.9 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26798]Diagnostic Scale - Transmit | Test: Test Name: | Diagnostic Scale - Transmit |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26797 - Run\_10-28\_11-8-25

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26797 |
| Configuration ID: | 14358 | Run Name: | Run\_10-28\_11-8-25 |
| Draft Run: | N | State: |  |
| Duration: | 1888 | Status: | Failed |
| Exec Date: | 28/10/16 | Target Cycle: |  |
| Exec Time: | 11:39:54 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_1.0.9 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26797]Diagnostic Scale - Record | Test: Test Name: | Diagnostic Scale - Record |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26795 - Run\_10-28\_11-5-43

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26795 |
| Configuration ID: | 14382 | Run Name: | Run\_10-28\_11-5-43 |
| Draft Run: | N | State: |  |
| Duration: | 131 | Status: | Passed |
| Exec Date: | 28/10/16 | Target Cycle: |  |
| Exec Time: | 11:07:55 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_1.0.9 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26795]Thermometer - Transmit | Test: Test Name: | Thermometer - Transmit |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26794 - Run\_10-28\_10-51-47

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26794 |
| Configuration ID: | 14381 | Run Name: | Run\_10-28\_10-51-47 |
| Draft Run: | N | State: |  |
| Duration: | 833 | Status: | Passed |
| Exec Date: | 28/10/16 | Target Cycle: |  |
| Exec Time: | 11:05:41 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_1.0.9 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26794]Thermometer - Record | Test: Test Name: | Thermometer - Record |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26793 - Run\_10-28\_10-46-7

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26793 |
| Configuration ID: | 14368 | Run Name: | Run\_10-28\_10-46-7 |
| Draft Run: | N | State: |  |
| Duration: | 330 | Status: | Passed |
| Exec Date: | 28/10/16 | Target Cycle: |  |
| Exec Time: | 10:51:37 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_1.0.9 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26793]Finger Oximeter - Transmit | Test: Test Name: | Finger Oximeter - Transmit |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26792 - Run\_10-28\_10-16-54

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26792 |
| Configuration ID: | 14353 | Run Name: | Run\_10-28\_10-16-54 |
| Draft Run: | N | State: |  |
| Duration: | 462 | Status: | Passed |
| Exec Date: | 28/10/16 | Target Cycle: |  |
| Exec Time: | 10:24:36 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_1.0.9 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26792]Finger Oximeter - Record | Test: Test Name: | Finger Oximeter - Record |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26791 - Run\_10-28\_9-47-0

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26791 |
| Configuration ID: | 14362 | Run Name: | Run\_10-28\_9-47-0 |
| Draft Run: | N | State: |  |
| Duration: | 1689 | Status: | Failed |
| Exec Date: | 28/10/16 | Target Cycle: |  |
| Exec Time: | 10:15:10 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_1.0.9 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26791]Tile Configuration | Test: Test Name: | Tile Configuration |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26790 - Run\_10-27\_15-48-1

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26790 |
| Configuration ID: | 14352 | Run Name: | Run\_10-27\_15-48-1 |
| Draft Run: | N | State: |  |
| Duration: | 263 | Status: | Failed |
| Exec Date: | 27/10/16 | Target Cycle: |  |
| Exec Time: | 15:52:25 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_1.0.9 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26790]Transmitter - General Screen | Test: Test Name: | Transmitter - General Screen |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26789 - Run\_10-27\_15-40-47

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26789 |
| Configuration ID: | 14351 | Run Name: | Run\_10-27\_15-40-47 |
| Draft Run: | N | State: |  |
| Duration: | 428 | Status: | Failed |
| Exec Date: | 27/10/16 | Target Cycle: |  |
| Exec Time: | 15:47:56 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_1.0.9 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26789]Clinician Messages - General Screen | Test: Test Name: | Clinician Messages - General Screen |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26787 - Run\_10-27\_15-35-2

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26787 |
| Configuration ID: | 14349 | Run Name: | Run\_10-27\_15-35-2 |
| Draft Run: | N | State: |  |
| Duration: | 165 | Status: | Passed |
| Exec Date: | 27/10/16 | Target Cycle: |  |
| Exec Time: | 15:37:48 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_1.0.9 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26787]Heart Failure Questionnaire | Test: Test Name: | Heart Failure Questionnaire |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26786 - Run\_10-27\_15-33-25

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26786 |
| Configuration ID: | 14348 | Run Name: | Run\_10-27\_15-33-25 |
| Draft Run: | N | State: |  |
| Duration: | 93 | Status: | Passed |
| Exec Date: | 27/10/16 | Target Cycle: |  |
| Exec Time: | 15:34:59 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_1.0.9 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26786]Blood Pressure - General Screen | Test: Test Name: | Blood Pressure - General Screen |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26785 - Run\_10-27\_15-28-53

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26785 |
| Configuration ID: | 14347 | Run Name: | Run\_10-27\_15-28-53 |
| Draft Run: | N | State: |  |
| Duration: | 233 | Status: | Passed |
| Exec Date: | 27/10/16 | Target Cycle: |  |
| Exec Time: | 15:32:46 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_1.0.9 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26785]Diagnostic Scale - General Screen | Test: Test Name: | Diagnostic Scale - General Screen |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26784 - Run\_10-27\_15-17-13

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26784 |
| Configuration ID: | 14380 | Run Name: | Run\_10-27\_15-17-13 |
| Draft Run: | N | State: |  |
| Duration: | 696 | Status: | Failed |
| Exec Date: | 27/10/16 | Target Cycle: |  |
| Exec Time: | 15:28:49 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_1.0.9 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26784]Thermometer - General Screen | Test: Test Name: | Thermometer - General Screen |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26783 - Run\_10-27\_15-14-18

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26783 |
| Configuration ID: | 14346 | Run Name: | Run\_10-27\_15-14-18 |
| Draft Run: | N | State: |  |
| Duration: | 145 | Status: | Failed |
| Exec Date: | 27/10/16 | Target Cycle: |  |
| Exec Time: | 15:16:44 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_1.0.9 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26783]Finger Oximeter - General Screen | Test: Test Name: | Finger Oximeter - General Screen |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26782 - Run\_10-27\_15-5-35

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26782 |
| Configuration ID: | 14336 | Run Name: | Run\_10-27\_15-5-35 |
| Draft Run: | N | State: |  |
| Duration: | 510 | Status: | Passed |
| Exec Date: | 27/10/16 | Target Cycle: |  |
| Exec Time: | 15:14:06 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_1.0.9 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26782]Logon - General and Patient Logon | Test: Test Name: | Logon - General and Patient Logon |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26780 - Run\_10-27\_14-49-42

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26780 |
| Configuration ID: | 14361 | Run Name: | Run\_10-27\_14-49-42 |
| Draft Run: | N | State: |  |
| Duration: | 825 | Status: | Failed |
| Exec Date: | 27/10/16 | Target Cycle: |  |
| Exec Time: | 14:55:15 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_1.0.9 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26780]Homepage - General Screen | Test: Test Name: | Homepage - General Screen |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26646 - Run\_10-17\_15-3-8

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26646 |
| Configuration ID: | 14369 | Run Name: | Run\_10-17\_15-3-8 |
| Draft Run: | N | State: |  |
| Duration: | 64 | Status: | Passed |
| Exec Date: | 17/10/16 | Target Cycle: |  |
| Exec Time: | 15:04:12 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi Clinical MediPiClinical\_v1.0.1-PILOT-20161006-1 / MediPiConcentrator\_v1.0.2-PILOT-20160921-1 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26646]Diagnostic Scale - Transmit | Test: Test Name: | Diagnostic Scale - Transmit |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26645 - Run\_10-17\_15-1-36

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26645 |
| Configuration ID: | 14370 | Run Name: | Run\_10-17\_15-1-36 |
| Draft Run: | N | State: |  |
| Duration: | 83 | Status: | Passed |
| Exec Date: | 17/10/16 | Target Cycle: |  |
| Exec Time: | 15:03:00 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi Clinical MediPiClinical\_v1.0.1-PILOT-20161006-1 / MediPiConcentrator\_v1.0.2-PILOT-20160921-1 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26645]Blood Pressure - Transmit | Test: Test Name: | Blood Pressure - Transmit |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26642 - Run\_10-17\_14-58-38

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26642 |
| Configuration ID: | 14368 | Run Name: | Run\_10-17\_14-58-38 |
| Draft Run: | N | State: |  |
| Duration: | 166 | Status: | Passed |
| Exec Date: | 17/10/16 | Target Cycle: |  |
| Exec Time: | 15:01:24 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi Clinical MediPiClinical\_v1.0.1-PILOT-20161006-1 / MediPiConcentrator\_v1.0.2-PILOT-20160921-1 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26642]Finger Oximeter - Transmit | Test: Test Name: | Finger Oximeter - Transmit |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26205 - Run\_9-29\_15-14-10

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26205 |
| Configuration ID: | 14368 | Run Name: | Run\_9-29\_15-14-10 |
| Draft Run: | N | State: |  |
| Duration: | 814 | Status: | Passed |
| Exec Date: | 29/09/16 | Target Cycle: |  |
| Exec Time: | 15:26:31 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.8a-PILOT-20160929-1 / MediPiConcentrator\_v1.0.2 / MediPiClinical\_v1.0.0 |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26205]Finger Oximeter - Transmit | Test: Test Name: | Finger Oximeter - Transmit |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26132 - Run\_9-23\_15-8-10

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26132 |
| Configuration ID: | 14366 | Run Name: | Run\_9-23\_15-8-10 |
| Draft Run: | N | State: |  |
| Duration: | 99 | Status: | Failed |
| Exec Date: | 23/09/16 | Target Cycle: |  |
| Exec Time: | 15:09:51 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.8-PILOT-20160907-1 (MediPi App/USB Connected Apps) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26132]Clinician Messages - Polling Warning | Test: Test Name: | Clinician Messages - Polling Warning |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26131 - Run\_9-23\_13-58-56

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26131 |
| Configuration ID: | 14365 | Run Name: | Run\_9-23\_13-58-56 |
| Draft Run: | N | State: |  |
| Duration: | 305 | Status: | Failed |
| Exec Date: | 23/09/16 | Target Cycle: |  |
| Exec Time: | 14:04:02 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.8-PILOT-20160907-1 (MediPi App/USB Connected Apps) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26131]MediPi Patient Banner - Patient Name | Test: Test Name: | MediPi Patient Banner - Patient Name |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26130 - Run\_9-23\_13-50-2

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26130 |
| Configuration ID: | 14364 | Run Name: | Run\_9-23\_13-50-2 |
| Draft Run: | N | State: |  |
| Duration: | 134 | Status: | Failed |
| Exec Date: | 23/09/16 | Target Cycle: |  |
| Exec Time: | 13:52:17 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.8-PILOT-20160907-1 (MediPi App/USB Connected Apps) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26130]Fatal Error Screen | Test: Test Name: | Fatal Error Screen |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26129 - Run\_9-23\_13-49-8

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26129 |
| Configuration ID: | 14363 | Run Name: | Run\_9-23\_13-49-8 |
| Draft Run: | N | State: |  |
| Duration: | 50 | Status: | Passed |
| Exec Date: | 23/09/16 | Target Cycle: |  |
| Exec Time: | 13:49:59 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.8-PILOT-20160907-1 (MediPi App/USB Connected Apps) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26129]Multiple Connected Devices | Test: Test Name: | Multiple Connected Devices |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26127 - Run\_9-23\_13-38-5

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26127 |
| Configuration ID: | 14362 | Run Name: | Run\_9-23\_13-38-5 |
| Draft Run: | N | State: |  |
| Duration: | 302 | Status: | Failed |
| Exec Date: | 23/09/16 | Target Cycle: |  |
| Exec Time: | 13:43:08 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.8-PILOT-20160907-1 (MediPi App/USB Connected Apps) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26127]Tile Configuration | Test: Test Name: | Tile Configuration |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26123 - Run\_9-23\_12-53-16

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26123 |
| Configuration ID: | 14352 | Run Name: | Run\_9-23\_12-53-16 |
| Draft Run: | N | State: |  |
| Duration: | 403 | Status: | Failed |
| Exec Date: | 23/09/16 | Target Cycle: |  |
| Exec Time: | 13:00:00 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.8-PILOT-20160907-1 (MediPi App/USB Connected Apps) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26123]Transmitter - General Screen | Test: Test Name: | Transmitter - General Screen |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26122 - Run\_9-23\_11-24-9

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26122 |
| Configuration ID: | 14351 | Run Name: | Run\_9-23\_11-24-9 |
| Draft Run: | N | State: |  |
| Duration: | 5335 | Status: | Not Completed |
| Exec Date: | 23/09/16 | Target Cycle: |  |
| Exec Time: | 12:53:05 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.8-PILOT-20160907-1 (MediPi App/USB Connected Apps) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26122]Clinician Messages - General Screen | Test: Test Name: | Clinician Messages - General Screen |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26121 - Run\_9-23\_10-57-24

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26121 |
| Configuration ID: | 14360 | Run Name: | Run\_9-23\_10-57-24 |
| Draft Run: | N | State: |  |
| Duration: | 484 | Status: | Failed |
| Exec Date: | 23/09/16 | Target Cycle: |  |
| Exec Time: | 11:00:00 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.8-PILOT-20160907-1 (MediPi App/USB Connected Apps) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26121]Scheduler - Transmit Summary Screen | Test: Test Name: | Scheduler - Transmit Summary Screen |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26120 - Run\_9-23\_10-47-34

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26120 |
| Configuration ID: | 14350 | Run Name: | Run\_9-23\_10-47-34 |
| Draft Run: | N | State: |  |
| Duration: | 550 | Status: | Passed |
| Exec Date: | 23/09/16 | Target Cycle: |  |
| Exec Time: | 10:56:45 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.8-PILOT-20160907-1 (MediPi App/USB Connected Apps) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26120]Scheduler - Schedule Screens | Test: Test Name: | Scheduler - Schedule Screens |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26119 - Run\_9-23\_10-42-16

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26119 |
| Configuration ID: | 14349 | Run Name: | Run\_9-23\_10-42-16 |
| Draft Run: | N | State: |  |
| Duration: | 282 | Status: | Passed |
| Exec Date: | 23/09/16 | Target Cycle: |  |
| Exec Time: | 10:46:58 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.8-PILOT-20160907-1 (MediPi App/USB Connected Apps) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26119]Heart Failure Questionnaire | Test: Test Name: | Heart Failure Questionnaire |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26118 - Fast\_Run\_9-23\_10-42-13

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: |  |
| Change Status: |  | Run ID: | 26118 |
| Configuration ID: | 14359 | Run Name: | Fast\_Run\_9-23\_10-42-13 |
| Draft Run: | N | State: |  |
| Duration: | 0 | Status: | Not Completed |
| Exec Date: | 23/09/16 | Target Cycle: |  |
| Exec Time: | 10:42:14 | Test Instance: |  |
| Host: |  | Test Set: Name: | MediPi\_v1.0.8-PILOT-20160907-1 (MediPi App/USB Connected Apps) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26118]Blood Pressure - Record | Test: Test Name: | Blood Pressure - Record |
| OS Build Number: |  | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: |  | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26117 - Run\_9-23\_10-33-31

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26117 |
| Configuration ID: | 14359 | Run Name: | Run\_9-23\_10-33-31 |
| Draft Run: | N | State: |  |
| Duration: | 379 | Status: | Failed |
| Exec Date: | 23/09/16 | Target Cycle: |  |
| Exec Time: | 10:39:50 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.8-PILOT-20160907-1 (MediPi App/USB Connected Apps) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26117]Blood Pressure - Record | Test: Test Name: | Blood Pressure - Record |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26116 - Run\_9-23\_10-30-25

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26116 |
| Configuration ID: | 14348 | Run Name: | Run\_9-23\_10-30-25 |
| Draft Run: | N | State: |  |
| Duration: | 244 | Status: | Passed |
| Exec Date: | 23/09/16 | Target Cycle: |  |
| Exec Time: | 10:33:25 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.8-PILOT-20160907-1 (MediPi App/USB Connected Apps) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26116]Blood Pressure - General Screen | Test: Test Name: | Blood Pressure - General Screen |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26115 - Run\_9-23\_10-0-13

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26115 |
| Configuration ID: | 14358 | Run Name: | Run\_9-23\_10-0-13 |
| Draft Run: | N | State: |  |
| Duration: | 1749 | Status: | Failed |
| Exec Date: | 23/09/16 | Target Cycle: |  |
| Exec Time: | 10:29:23 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.8-PILOT-20160907-1 (MediPi App/USB Connected Apps) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26115]Diagnostic Scale - Record | Test: Test Name: | Diagnostic Scale - Record |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26114 - Run\_9-23\_9-57-13

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26114 |
| Configuration ID: | 14347 | Run Name: | Run\_9-23\_9-57-13 |
| Draft Run: | N | State: |  |
| Duration: | 174 | Status: | Passed |
| Exec Date: | 23/09/16 | Target Cycle: |  |
| Exec Time: | 10:00:07 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.8-PILOT-20160907-1 (MediPi App/USB Connected Apps) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26114]Diagnostic Scale - General Screen | Test: Test Name: | Diagnostic Scale - General Screen |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26113 - Run\_9-23\_9-29-12

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26113 |
| Configuration ID: | 14353 | Run Name: | Run\_9-23\_9-29-12 |
| Draft Run: | N | State: |  |
| Duration: | 1522 | Status: | Failed |
| Exec Date: | 23/09/16 | Target Cycle: |  |
| Exec Time: | 09:54:34 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.8-PILOT-20160907-1 (MediPi App/USB Connected Apps) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26113]Finger Oximeter - Record | Test: Test Name: | Finger Oximeter - Record |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26112 - Run\_9-23\_9-25-24

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26112 |
| Configuration ID: | 14346 | Run Name: | Run\_9-23\_9-25-24 |
| Draft Run: | N | State: |  |
| Duration: | 223 | Status: | Failed |
| Exec Date: | 23/09/16 | Target Cycle: |  |
| Exec Time: | 09:29:07 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.8-PILOT-20160907-1 (MediPi App/USB Connected Apps) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26112]Finger Oximeter - General Screen | Test: Test Name: | Finger Oximeter - General Screen |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26111 - Run\_9-23\_9-7-54

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26111 |
| Configuration ID: | 14344 | Run Name: | Run\_9-23\_9-7-54 |
| Draft Run: | N | State: |  |
| Duration: | 1033 | Status: | Passed |
| Exec Date: | 23/09/16 | Target Cycle: |  |
| Exec Time: | 09:25:08 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.8-PILOT-20160907-1 (MediPi App/USB Connected Apps) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26111]Logon - MediPi Session Lock | Test: Test Name: | Logon - MediPi Session Lock |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26108 - Run\_9-22\_14-18-12

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26108 |
| Configuration ID: | 14361 | Run Name: | Run\_9-22\_14-18-12 |
| Draft Run: | N | State: |  |
| Duration: | 493 | Status: | Passed |
| Exec Date: | 22/09/16 | Target Cycle: |  |
| Exec Time: | 16:00:16 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.8-PILOT-20160907-1 (MediPi App/USB Connected Apps) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26108]Homepage - General Screen | Test: Test Name: | Homepage - General Screen |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |

## Run ID : 26096 - Run\_9-20\_14-41-8

|  |  |  |  |
| --- | --- | --- | --- |
| Field Label | Field Value | Field Label | Field Value |
| Baseline: |  | Operating System: | Windows 7 |
| Change Status: |  | Run ID: | 26096 |
| Configuration ID: | 14336 | Run Name: | Run\_9-20\_14-41-8 |
| Draft Run: | N | State: |  |
| Duration: | 1106 | Status: | Passed |
| Exec Date: | 20/09/16 | Target Cycle: |  |
| Exec Time: | 15:02:02 | Test Instance: |  |
| Host: | 224610-WIN7 | Test Set: Name: | MediPi\_v1.0.8-PILOT-20160907-1 (MediPi App/USB Connected Apps) |
| Iterations Status: |  | Test Version Number: |  |
| Test Configuration: Name: | [26096]Logon - General and Patient Logon | Test: Test Name: | Logon - General and Patient Logon |
| OS Build Number: | Build 7601 | Tester: | stwi2(Stephen Willis) |
| OS Service Pack: | Service Pack 1 | Type: | MANUAL |

|  |
| --- |
| Comments |
|  |